Design Verification Plan

for

Advanced Visualization Workspace (AVW) 15,

CT ASPECT Scoring (CT ASPECTS) & NM Suite

Doc ID: D000995566

Document Revision: H

1	Document Introduction	3

1.1	Purpose	3

1.2	Scope	4

1.3	References	10

1.4	Abbreviations and Definitions	16

2	Strategy	17

3	Test Design	17

3.1	Test Protocol	17

3.2	Verification Activities Approach Overview	20

3.3	Methods Overview	21

3.4	Tools Overview	31

3.5	Test Environment Overview	31

3.6	Product Hazard/Risk Management Activities	45

3.7	Exploratory Testing Strategy	48

3.8	Interface Testing Strategy	48

3.9	Usability Engineering - Formative Evaluation Strategy	49

4	Resources	50

4.2	Capital Equipment	56

4.3	Key Personnel	58

4.4	Estimates	60

5	Verification Samples	63

6	Entry Criteria	63

7	Exit Criteria	64

7.1	Suspension Criteria:	64

8	Acceptance Criteria	64

9	Defect Management	64

10	Record history	65

11	Record Review and Approval	66

12	Appendices	66

12.1	Management of Verification Activities	66

12.2	Verification Table (V-Table)	70

12.3	Exploratory Testing	76

12.4	Expert Testing	77

# Document Introduction

## Purpose

This document plans (not the scheduling) the management, execution, and control of the verification activities of the requirements in scope of this document.

The objective of the Verification Plan document is to plan (i.e., Identify the activities, methods, resources, etc.) the comprehensive verification activities that:

- Discover defects as early as possible in the product lifecycle.

- Assure conformance to standards.

- Increase confidence level in the reliability of the product.

- Provide objective evidence that the “product was built right”.

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

The purpose of this plan is to identify the tasks, deliverables and resources required for Verification of ‘ISP13 VAL-2’ release.

RMM Tests Execution in ISP13 VAL-3 Development Phase: This document is an extension of VAL-2 released Verification Plan (Rev A) for enabling execution of RMM tests during Development Phase of ISP13 VAL3 on VAL3 released Verification ISOs.

#### ISP13 VAL-3 Release

The purpose of this plan is to identify the tasks, deliverables and resources required for Verification of ‘ISP13 VAL-3’ release.

RMM Tests Execution in AVW15 RFD Development Phase: This document is an extension of ISP13 VAL-3 released Verification Plan (Rev D) for enabling execution of RMM tests during Development Phase of ISP13 VAL3 on VAL3 released Verification ISOs.

#### AVW15 RFD Release

The purpose of this plan is to identify the tasks, deliverables and resources required for Verification of ‘AVW15 RFD’ release.

#### 1.1.1.4 AVW15 RC Release

The purpose of this plan is to identify the tasks, deliverables and resources required for Verification of ‘AVW15 RC’ release.

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

The purpose of this plan is to identify the tasks, deliverables and resources required for Verification of ‘CT ASPECTS’ RFD release.

#### 1.1.2.2 RC Release

The purpose of this plan is to identify the tasks, deliverables and resources required for Verification of ‘CT ASPECTS’ RC release.

### NM Suite

#### RFD Release

The purpose of this plan is to identify the tasks, deliverables and resources required for Verification of ‘NM Suite’ RFD release.

#### RC Release

#### The purpose of this plan is to identify the tasks, deliverables and resources required for Verification of ‘NM Suite’ RC release.

## Scope

The scope of this plan includes Advanced Visualization workspace V15, previously named IntelliSpace Portal V13, CT ASPECTS V15 and NM Suite V15. NM suite and CT ASPECTS are launched from AVW environment and are part of its installation package. Therefor all the products versions will be identified under the same version number (V15).

### Advance Visualization Workspace 15

Advance Visualization Workspace V15, previously named as IntelliSpace Portal V13 (ISP13) introduces a new version of the Advance Visualization product. The purpose of this document is to list the planned tests on AVW15 for different releases.

#### ISP13 VAL-2 Release

Below is the scope of testing for ISP13 VAL-2 release. Refer section 12.2 for more information. The details of the features that will be tested are listed in section 12.2.2.

##### MR Cardiac Suite: Full Functionality & RMM tests execution for below:

  - Viewer

  - Functional LV & RV

  - QFlow

  - Spatial Enhancement

  - Functional LA

##### CT Applications:

- CT Liver – Full Functionality (DRS Coverage) tests execution

- CCA - Full Functionality tests of ISP13 features including regression & RMF tests execution.

##### RMM Tests Execution in ISP13 VAL-3 Development Phase for non-impacted

- MR Apps

- NM Apps

- CT Apps

- CT Cardiac Apps

- CT Spectral Apps

- CT Viewing

#### ISP13 VAL-3 Release

Below is the scope of testing for ISP13 VAL-3 release. Refer section 12.2 for more information. The details of the features that will be tested are listed in section 12.2.2.

##### MR Cardiac Suite:

- Temporal Analysis

- Functional LV RV

- Mapping

##### MR Applications:

- LOBI CVS

##### CT Applications:

- sAVA

- Auto CS

- Trauma

- AVA

##### CT Spectral Applications:

- sMGOP/LMG 3D and MPR

- sCTV RFCs

##### RMM Tests Execution for Legacy:

- MR Cardiac Applications

- MR Applications

- CT Applications

- NM Applications

##### System Features:

- NLog

- Regression

- RMM

- PMCF for MR cardiac

- Labeling

##### RMM Tests Execution in AVW15 RFD Development Phase for non-impacted

- MR Apps

- CT Apps

- CT Cardiac Apps

- CT Spectral Apps

- CT Viewing

- NM Apps

- CT, MR, NM Legacy Apps

- System Features

#### AVW15 RFD Release

Below is the scope of testing for AVW15 RFD release. Refer section 12.2 for more information. The details of the features that will be tested are listed in section 12.2.2.

##### MR Cardiac Suite:

- Findings, Bookmarks & Save Functionality

- RMM Tests

- Regression Tests

##### MR Applications:

- RMM Tests

##### CT Applications:

- Standard Reporting

- RMM Tests

- Regression Tests

###### Updates based on the change in Test Strategy for AVW15 RFD Verification

- Below mentioned PRS Keys of Liver Analysis application are not in scope:

    - 1283709: ISP_PRS_ID_1_3: Liver Analysis Application

    - 1283710: ISP_PRS_ID_1_4: Liver Analysis Application veins segmentation

    - 1283711: ISP_PRS_ID_1_7: Liver Analysis Application

- In Scope – Impacted tests of Liver Analysis application

##### System Features HFA:

- ZFP

- Labeling

- 3rd Party application on Client

- PMCF Updated

- Regression Tests

##### CT Cardiac Applications:

- CAD RADS and Findings Management

- Discrete Report – CCA and TAVI

- CCA and TAVI Legacy

- RMM Tests

##### CT Spectral Applications:

- sMGOP/LMG 3D and 2D support

- sCTV RFCs

- RMM Tests

##### CT Viewing:

- RMM Tests

##### NM Applications:

- 3rd Party Applications

    - Regression Tests

##### CT, MR, NM Legacy Applications:

- RMM Tests

###### Updates based on the change in Test Strategy for AVW15 RFD Verification

For PRS Key “529576: ISPV7_PRS_ID_5_46: NM Compatibility EBW-NM” older version of released IntelliSpace Portal (ISP12) will be used for testing & comparison instead of EBW-NM.

Note: 1. EBW-NM runs on Windows XP which is obsolete.
          2. ISP12 to EBW-NM compatibility is established and verified. Hence, ISP12 will be        
               used as base for this comparison.

##### Installation, Upgrades & Security:

- 3rd Party Decoupling

- Client Installer Re-design

- Regression Tests

##### System Features BLR:

- DynaCad 5.0

- Regression Tests

- RMM Tests

##### Serviceability:

- IOT

- Regression Tests

##### Non-Functional Tests:

- Performance Tests of Applications

##### Compatibility Tests:

| ## | Modality | AVW15 RFD |  |
| --- | --- | --- | --- |
| 1 | CT | CT HAWK v5.0 | BB v4.8 |
| 2 | MR | R5.9 |  |
| 3 | NM | Vereos Digital PET/CT HeartBeat V2.0.3 | Legacy NM cameras - Forte JETStream |
| 4 | ISCV | V6.2 |  |
| 5 | IBE | B.18 |  |
| 6 | ISPACS | ISPACS 4.4.553 | ISPACS ISR 4.7.11 UDM 3.1.8 |
| 7 | VuePACS | VuePACS 12.2.8 |  |
| 8 | DDD | CardioMD IV - legacy |  |
| 9 | US | Affiniti 50C, Affiniti 70C, EPIQ 5C, EPIQ 7C, EPIQ CVx_9 |  |

#### AVW15 RC Release

##### MR Cardiac Suite:

- Functional Test

- RMM

##### MR Applications:

- Sanity and Regression Tests (including LOBI and NQ changes)

##### CT Applications:

- Regression Tests

- BP update regression

- AVA vessels tool functional test

##### System Features HFA:

- Labeling and Regression Tests

##### CT Cardiac Applications:

- CAD RADS 2.0 and Findings Management regression test

- CCA RFCs regression test

##### CT Spectral Applications:

- Regression

- sCTV and sMGOP new requirement Functionality Test

- RMM (impact based)

##### CT Viewing:

- Regression test, Sanity and Upgrades

- CT Viewing - Angio CT regression test

##### NM Applications:

- NM Suite Sanity Testing

- NM 3rd Party Apps Sanity & Regression (Autoquant 2017.22)

- MiradaXD Integration Sanity & Regression

- Knowledgescape Sanity & Regression

##### CT, MR, NM Legacy Applications:

- MR RMM (impact based)

- CT, NM (including Corridor4DM 2023) MRC, and MR Regression Tests

##### Installation, Upgrades & Security:

- Installation, upgrades, and Security Regression Tests

- PIAB with both C and D drive SSD regression

- Enable Encryption option on PMT.

##### System Features BLR:

- PMT client auto upgrade Functional tests.

- Regression Tests

- DynaCAD V5.1

##### Serviceability:

- IOT RTS1.0  Functional test

- PCS 9.0.1-148 Regression Test

- Axeda Split package and SPLUNK Regression test

- Regression Tests

##### Non-Functional Tests:

- Performance Tests of Applications

- Reliability Test

##### ISP Sanity:

  - Gen3 and A2 GPU

##### Compatibility Tests:

| # | Modality | ISP13 RC |
| --- | --- | --- |
| 1 | CT | CT Incisive patch (couch height fix) Incisive 1.1 HB, V4.5.1.742 |
| 2 | CT | CT Incisive for Dual Energy verification per marketing request |
| 2 | CT | CT Thor (if the preliminary datacases are available on time) |
| 3 | ISCV | V7.0 |
| 4 | IBE | B.19 (if available on time) |
| 5 | iGT (IXR) | Allura 8.x.100 |
| 6 | US | More legacy cameras if the data is available |
| 7 | Pinnacle | IOCC verification only |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

- Basic Functionality

#### RC Release

- Sanity

### NM Suite

#### RFD Release

- NM Planar Suite

- NM JETPack

- NM QA QC Tools

Note: There is no change from the previous NM Suite product and hence, only RMM and Regression tests are planned for execution.

#### RC Release

- Sanity for NM Planar Suite, NM JETPack and NM QA QC Tools

## References

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

| Reference | Document Title | Document ID | Revision |
| --- | --- | --- | --- |
| REF – 1 | ISP13 Product Release Plan | DHF380218 | LB |
| REF – 2 | ISP13 User Requirement specification | DHF380082 | LD |
| REF – 3 | ISP13 Product Requirement specification | DHF380211 | LC |
| REF – 4 | ISP13 Usability Engineering Plan | DHF380228 | LB |
| REF – 5 | ISP13 Labeling Plan | DHF381090 | LB |
| REF – 6 | ISP13 Label Specification | D000941165 | B |
| REF – 7 | ISP13 System Design Specifications | DHF380234 | LB |
| REF – 8 | ISP13 Risk Management Plan | DHF380156 | LB |
| REF – 9 | ISP13 Reliability plan | D000933496 | A |
| REF – 10 | ISP13 Product Security Risk Management Report | DHF381138 | LA |
| REF – 11 | Tool List for ISP Product line | DHF379868 | LD |
| REF – 12 | ISP13 3rd Party items assessment checklist | DHF380838 | A |
| REF – 13 | ISP13 Cardiac suite DRS | D000856146 | A |
| REF – 14 | ISP13 MR Cardiac suite Viewer Verification Spec | D000994617 | A |
| REF – 15 | ISP13 MR Cardiac suite Functional LVRV Verification Spec | D000996525 | A |
| REF – 16 | ISP13 MR Cardiac suite QFlow Verification Spec | D000996541 | A |
| REF – 17 | ISP13 MR Cardiac suite Functional Long Axis Verification Spec | D000996542 | A |
| REF – 18 | ISP13 RMA MR Cardiac Suite | D000991970 | A |
| REF – 19 | ISP13 Risk Management Matrix Cardiac Suite | D000879224 | A |
| REF – 20 | ISP13 MR Cardiac Suite Traceability Matrix | D000996539 | A |
| REF – 21 | ISP13 DRS CT Liver | D000820457 | A |
| REF – 22 | ISP13 Test Specification HFA (CT Liver) | D000995159 | A |
| REF – 23 | ISP13 RMA for CTApps | D000938106 | C |
| REF – 24 | ISP13 RMM General CTApps | D000892550 | C |
| REF – 25 | ISP13 Traceability Matrix CTApps | D000993523 | A |
| REF – 26 | ISP13 CCA DRS | D000872156 | B |
| REF – 27 | ISP13 DRS _TAVI | D000865239 | A |
| REF – 28 | ISP13 Test Specifications CT Cardiac | D000946877 | A |
| REF – 29 | ISP13 Sanity Spec CT Cardiac | D000998132 | A |
| REF – 30 | ISP13 RMA CT Cardiac | D000946872 | A |
| REF – 31 | ISP13_Risk Management Matrix_CT Cardiac | D000946860 | A |
| REF – 32 | ISP13 Traceability Matrix CT Cardiac | D000946874 | A |

RMM Tests Execution in ISP13 VAL-3 Development Phase:

| Reference | Document Title | Document ID | Revision |
| --- | --- | --- | --- |
| REF – 33 | ISP13_PVT_TestSpec | D001058943 | A |
| REF – 34 | ISP13_Risk Management Analysis_MRApps & MMV | D001065109 | A |
| REF – 35 | ISP13_RMM_MRAPPS & MMV | D001065108 | A |
| REF – 36 | ISP13 Legacy MR Cardiac RMA | D001065179 | A |
| REF – 37 | ISP13 MR Cardiac Legacy applications Hazard Test Spec | D001064798 | A |
| REF – 38 | ISP13 NM Apps_RMA | D001059457 | A |
| REF – 39 | ISP13 NM Apps_RMM | D001059463 | A |
| REF – 40 | ISP13 NM Apps_Sanity Test Specification | D001059465 | A |
| REF – 41 | ISP13 RMA for CTApps | D000938106 | C |
| REF – 42 | ISP13 RMM General CTApps | D000892550 | C |
| REF – 43 | ISP13 Test Specification HFA | D000995159 | C |
| REF – 44 | ISP13 RMA CT Cardiac | D000946872 | A |
| REF – 45 | ISP13_Risk Management Matrix_CT Cardiac | D000946860 | A |
| REF – 46 | ISP13 Test Specifications CT Cardiac | D000946877 | D |
| REF – 47 | ISP13 DRS sMGoP | D000882706 | B |
| REF – 48 | ISP13 DRS General Spectral Viewing | D000822200 | A |
| REF – 49 | ISP13 RMA CTSpectral Apps | D000952406 | A |
| REF – 50 | ISP13 RMM CTSpectral Apps | D000996304 | A |
| REF – 51 | ISP13 Test spec CTSpectral sMG | D001046073 | A |
| REF – 52 | ISP13 CTV Functional Regression Test Specification | D000999069 | B |

#### ISP13 VAL-3 Release

| Reference | Document Title | Document ID | Revision |
| --- | --- | --- | --- |
| REF – 1 | ISP13 Product Release Plan | DHF380218 | LC |
| REF – 2 | ISP13 User Requirement specification | DHF380082 | LD |
| REF – 3 | ISP13 Product Requirement specification | DHF380211 | LD |
| REF – 4 | ISP13 Usability Engineering Plan | DHF380228 | LB |
| REF – 5 | ISP13 Labeling Plan | DHF381090 | LB |
| REF – 6 | ISP13 Label Specification | D000941165 | B |
| REF – 7 | ISP13 System Design Specifications | DHF380234 | LB |
| REF – 8 | ISP13 Risk Management Plan | DHF380156 | LB |
| REF – 9 | ISP13 Reliability plan | D000933496 | A |
| REF – 10 | ISP13 Product Security Risk Management Report | DHF381138 | LA |
| REF – 11 | Tool List for ISP Product line | DHF379868 | LD |
| REF – 12 | ISP13 3rd Party items assessment checklist | DHF380838 | A |
| REF – 13 | ISP13 Cardiac suite DRS | D000856146 | C |
| REF – 14 | ISP13 LOBI DRS | D001036765 | A |
| REF – 15 | ISP13_TestSpecMR Cardiac Functional LV RV | D000996525 | A |
| REF – 16 | ISP13 Temporal Enhancement Verification Specification | D001104277 | A |
| REF – 17 | ISP13 MR Cardiac Suite Mapping Functionality Test specification | D001104051 | A |
| REF – 18 | ISP13 LoBI Regression Verification Spec | D001104696 | A |
| REF – 19 | ISP13 RMA MR Cardiac Suite | D000991970 | A |
| REF – 20 | ISP13 Risk Management Matrix Cardiac Suite | D000879224 | B |
| REF – 21 | ISP13 LoBI RMA | D001105931 | A |
| REF – 22 | ISP13 LoBI RMM | D001104871 | A |
| REF – 23 | ISP13_MR Cardiac Suite Traceability Matrix | D000996539 | B |
| REF – 24 | ISP13 Traceability Matrix LoBI | D001105932 | A |
| REF – 25 | ISP13 NM Apps_Sanity Test Specification | D001059465 | B |
| REF – 26 | ISP13 NM Apps RMF Test Specification TRR | D001107914 | A |
| REF – 27 | ISP13 NM Apps_RMA | D001059457 | A |
| REF – 28 | ISP13 NM Apps_RMM | D001059463 | A |
| REF – 29 | ISP13_Verification Test Specification and Report CT Cardiac | D000946877 | D |
| REF – 30 | ISP13 RMA CT Cardiac | D000946872 | A |
| REF – 31 | ISP13_Risk Management Matrix_CT Cardiac | D000946860 | A |
| REF – 32 | ISP13 DRS AVA | D001083895 | A |
| REF – 33 | ISP13 DRS Automated Calcium Scoring Results | D001012896 | A |
| REF – 34 | Test Specification and Report HFAAPPS | D000995159 | D |
| REF – 35 | ISP13 RMA for CTapps | D000938106 | D |
| REF – 36 | ISP13 RMM General CTApps | D000892550 | D |
| REF – 37 | Traceability Matrix CT Apps | D000993523 | B |
| REF – 38 | ISP13 SMGOP DRS | D000882706 | C |
| REF – 39 | ISP13 DRS Spectral CT Viewer | D001085961 | A |
| REF – 40 | ISP13 Spectral General Viewing DRS | D000822200 | C |
| REF – 41 | ISP13 Study approved CTSpectral apps | D001104395 | A |
| REF – 42 | ISP13 Test Specification-Spectral CT Multiphase Analysis | D001093929 | A |
| REF – 43 | ISP13 Test Spec CT Spectral sCTV | D001092392 | A |
| REF – 44 | ISP13 Test spec CTSpectral sMG | D001046073 | B |
| REF – 45 | ISP13 Test Script Validation Report-CTSpectral Automation | D001047485 | B |
| REF – 46 | ISP13 RMA CTSpectral Apps | D000952406 | B |
| REF – 47 | ISP13 RMM CTSpectral Apps | D000996304 | B |
| REF – 48 | ISP13 Traceability Matrix for CTSpectral Applications | D001097887 | A |
| REF – 49 | ISP13_PVT_TestSpec | D001058943 | A |
| REF – 50 | ISP13_TestSpecification_SystemFeatures-BLR | D000999142 | A |
| REF – 51 | ISP13 System RMM | D000996286 | A |

RMM Tests Execution in AVW15 RFD Development Phase:

| Reference | Document Title | Document ID | Revision |
| --- | --- | --- | --- |
| REF – 1 | ISP13 RMA MRCardiac Suite | D000991970 | B |
| REF – 2 | ISP13 RISK MANAGEMENT MATRIX CARDIAC SUITE | D000879224 | C |
| REF – 3 | ISP13 MRCardiac suite Functional LVRV Verification spec | D000996525 | B |
| REF – 4 | ISP13 MR Cardiac Suite-QFlow Verification Specification | D000996541 | B |
| REF – 5 | ISP13 MR Cardiac Suite-Functional Long Axis Verification Specification | D000996542 | B |
| REF – 6 | ISP13 MR Cardiac Suite Unit/Sanity Test Specification | D001001801 | C |
| REF – 7 | ISP13_CardiacMapping_Test Specification | D001104051 | A |
| REF – 8 | ISP13 Temporal Enhancement Test Specification and Report | D001104277 | A |
| REF – 9 | ISP13 MR Cardiac Suite-Spatial Enhancement Verification Spec | D000996623 | B |
| REF – 10 | ISP13 MR Cardiac Hazard Verification Spec Legacy Application | D001064798 | A |
| REF – 11 | ISP13 MR Cardiac suite Viewer Verification Spec | D000994617 | C |
| REF – 12 | ISP13_Risk Management Analysis_MRApps & MMV | D001065109 | A |
| REF – 13 | ISP13 RMA LoBI | D001105931 | A |
| REF – 14 | Risk Management Matrix for ISP13,MR Applications and MMV | D001104871 | B |
| REF – 15 | ISP13_RMM_MRAPPS & MMV | D001065108 | A |
| REF – 16 | ISP13 Regression Spec LOBI | D001104696 | A |
| REF – 17 | ISP13 RMA for CTapps | D000938106 | D |
| REF – 18 | ISP13 RMM General CTApps | D000892550 | D |
| REF – 19 | Test Specification HFA | D000995159 | E |
| REF – 20 | ISP13 Automation Validation Report for CT Apps and BCKHFA | D000996313 | A |
| REF – 21 | ISP13 RMA CT Cardiac | D000946872 | A |
| REF – 22 | ISP13 RMM CT Cardiac | D000946860 | A |
| REF – 23 | ISP13 Sanity Spec CT Cardiac | D000998132 | A |
| REF – 24 | ISP13 Test Specifications CT Cardiac | D000946877 | E |
| REF – 25 | ISP13 RMA CTSpectral Apps | D000952406 | B |
| REF – 26 | ISP13 RMM CTSpectral Apps | D000996304 | B |
| REF – 27 | ISP13 Study approved CTSpectral apps | D001104395 | A |
| REF – 28 | ISP13 Test Script Validation Report-CTSpectral Automation | D001047485 | C |
| REF – 29 | Test Specification-Spectral CT Multiphase Analysis | D001093929 | B |
| REF – 30 | ISP13 Test Spec CT Spectral sCTV | D001092392 | E |
| REF – 31 | ISP13 Test spec CTSpectral sMG | D001046073 | E |
| REF – 32 | ISP13 CTV Functional Regression Test Specification | D000999069 | B |
| REF – 33 | ISP13 NM Apps_RMA | D001059457 | A |
| REF – 34 | ISP13 NM Apps_RMM | D001059463 | A |
| REF – 35 | ISP13 NM Apps_Sanity Test Specification | D001059465 | B |
| REF – 36 | ISP13 NM Apps RMF Test Specification TRR | D001107914 | A |
| REF – 37 | ISP13 PVT Owned CT Apps sanity Test Spec | D001048147 | A |
| REF – 38 | ISP13_PVT_TestSpec | D001058943 | A |
| REF – 39 | ISP13 RMA for ISP System | DHF380158 | LC |
| REF – 40 | ISP13 System RMM | D000996286 | A |
| REF – 41 | ISP13_TestSpecification_SystemFeatures-BLR | D000999142 | B |

#### AVW15 RFD Release

| Reference | Document Title | Document ID | Revision |
| --- | --- | --- | --- |
| REF – 1 | ISP13 Product Release Plan | DHF380218 | LE |
| REF – 2 | ISP13 User Requirement specification | DHF380082 | LE |
| REF – 3 | ISP13 Product Requirement specification | DHF380211 | LG |
| REF – 4 | ISP13 Usability Engineering Plan | DHF380228 | LC |
| REF – 5 | ISP13 Labeling Plan | DHF381090 | LC |
| REF – 6 | ISP13 Label Specification | D000941165 | D |
| REF – 7 | ISP13 System Design Specifications | DHF380234 | LD |
| REF – 8 | ISP13 Risk Management Plan | DHF380156 | LD |
| REF – 9 | CT ASPECTS Design Verification Plan | D001157489 | A |
| REF – 10 | ISP13 Reliability plan | D000933496 | B |
| REF – 11 | ISP13 Product Security Risk Management Report | DHF381138 | LC |
| REF – 12 | Tool List for ISP Product line | DHF379868 | LE |
| REF – 13 | ISP13 3rd Party items assessment checklist | DHF380838 | C |
| REF – 14 | ISP13 Compatibility Impact Analysis | D000819409 | C |
| REF – 15 | ISP13 URS-PRS Traceability | D001105259 | A |
| REF – 16 | ISP13 Cardiac suite DRS | D000856146 | E |
| REF – 17 | ISP13_TestSpecMR Cardiac Functional LV RV | D000996525 | C |
| REF – 18 | ISP13 Temporal Enhancement Verification Specification | D001104277 | C |
| REF – 19 | ISP13_CardiacMapping_Test Specification | D001104051 | C |
| REF – 20 | ISP13_TestSpecMRCardiacQflow | D000996541 | C |
| REF – 21 | ISP13_TestSpecMR Cardiac Functional Long Axis | D000996542 | C |
| REF – 22 | ISP13_MR Cardiac Suite Sanity_Unit Test Specification | D001001801 | F |
| REF – 23 | ISP13 MR Cardiac Suite-Spatial Enhancement Verification Spec | D000996623 | C |
| REF – 24 | ISP13 MR Cardiac suite Viewer Verification Spec | D000994617 | E |
| REF – 25 | ISP13 RMA MR Cardiac Suite | D000991970 | C |
| REF – 26 | ISP13 Risk Management Matrix Cardiac Suite | D000879224 | D |
| REF – 27 | ISP13_MR Cardiac Suite Traceability Matrix | D000996539 | E |
| REF – 28 | ISP13_Risk Management Analysis_MRApps & MMV | D001065109 | A |
| REF – 29 | ISP13_RMM_MRAPPS & MMV | D001065108 | A |
| REF – 30 | Test Specification and Report HFAAPPS | D000995159 | F |
| REF – 31 | Traceability Matrix CT Apps | D000993523 | E |
| REF – 32 | ISP13 RMA for CTapps | D000938106 | E |
| REF – 33 | ISP13 RMM General CTApps | D000892550 | E |
| REF – 34 | Test Specification and Report SystemHFA | D000996325 | B |
| REF – 35 | Traceability Matrix System HFA | D001159554 | A |
| REF – 36 | ISP13 CCA DRS | D000872156 | C |
| REF – 37 | ISP13 TAVI DRS | D000865239 | B |
| REF – 38 | ISP13_Verification Test Specification and Report CT Cardiac | D000946877 | G |
| REF – 39 | ISP13_Sanity Test Specification and Report CT Cardiac | D000998132 | A |
| REF – 40 | ISP13 Traceability Matrix CT Cardiac | D000946874 | E |
| REF – 41 | ISP13_Risk Management Matrix_CT Cardiac | D000946860 | B |
| REF – 42 | ISP13 RMA CT Cardiac | D000946872 | B |
| REF – 43 | ISP13 SMGOP DRS | D000882706 | D |
| REF – 44 | ISP13 DRS Spectral CT Viewer | D001085961 | B |
| REF – 45 | ISP13 Spectral General Viewing DRS | D000822200 | D |
| REF – 46 | AVW15 DRS Spectral CTMA | D001142845 | A |
| REF – 47 | Test Specification-Spectral CT Multiphase Analysis | D001093929 | D |
| REF – 48 | ISP13 Test Spec CT Spectral sCTV | D001092392 | H |
| REF – 49 | ISP13 Test spec CTSpectral sMG | D001046073 | G |
| REF – 50 | ISP13 Test Script Validation Report-CTSpectral Automation | D001047485 | D |
| REF – 51 | ISP13-Code TRR CTSpectral Automation | D001047484 | D |
| REF – 52 | ISP13 Study approved CTSpectral apps | D001104395 | A |
| REF – 53 | ISP13 Traceability Matrix for CTSpectral Applications | D001097887 | C |
| REF – 54 | ISP13 RMA CTSpectral Apps | D000952406 | C |
| REF – 55 | ISP13 RMM CTSpectral Apps | D000996304 | C |
| REF – 56 | ISP13 CTV Functional Regression Test Specification | D000999069 | C |
| REF – 57 | ISP13 NM Apps_Sanity Test Specification | D001059465 | B |
| REF – 58 | ISP13 NM Apps RMF Test Specification TRR | D001107914 | B |
| REF – 59 | AVW15_DRS_System_Installation | D001149811 | A |
| REF – 60 | ISP13_PVT_Installation & Upgrades_TestSpec | D001115079 | D |
| REF – 61 | ISP13 PVT Owned CT Apps sanity Test Spec | D001048147 | A |
| REF – 62 | ISP13_PVT_TestSpec | D001058943 | B |
| REF – 63 | PVT_AVW15_Installation and Upgrades_Traceability | D001157062 | A |
| REF – 64 | DRS for Remote Tools Suite Installation & Features Requirements for AVW / Console / Enterprise | D001137317 | A |
| REF – 65 | ISP13_TestSpecification_SystemFeatures-BLR | D000999142 | D |
| REF – 66 | AVW15_TestSpecificion_Serviceability | D001164936 | A |
| REF – 67 | AVW15_Traceability_SystemFeatures-BLR & Serviceability | D001164937 | A |
| REF – 68 | ISP13 System RMM | D000996286 | B |
| REF – 69 | ISP13 DRS System Performance | D001157240 | A |
| REF – 70 | AVW15 NFT Performance Test Specification | D001163851 | A |
| REF – 71 | AVW15_Consolidated Code TRR for NFTPerformanceAutomation | D001163854 | A |
| REF – 72 | AVW15_Automation Test Script Validation Report_NFT_Performance | D001163852 | A |
| REF – 73 | AVW15 Traceability Matrix_NFT | D001163855 | A |

#### AVW15 RC Release

| Reference | Document Title | Document ID | Revision |
| --- | --- | --- | --- |
| REF – 1 | ISP13 Product Release Plan | DHF380218 | LF |
| REF – 2 | ISP13 User Requirement specification | DHF380082 | LG |
| REF – 3 | ISP13 Product Requirement specification | DHF380211 | LI |
| REF – 4 | ISP13 Usability Engineering Plan | DHF380228 | LE |
| REF – 5 | ISP13 Labeling Plan | DHF381090 | LE |
| REF – 6 | ISP13 Label Specification | D000941165 | E |
| REF – 7 | ISP13 System Design Specifications | DHF380234 | LE |
| REF – 8 | ISP13 Risk Management Plan | DHF380156 | LF |
| REF – 9 | CT ASPECTS Design Verification Plan | D001157489 | A |
| REF – 10 | ISP13 Reliability plan | D000933496 | C |
| REF – 11 | ISP13 Product Security Risk Management Report | DHF381138 | LE |
| REF – 12 | Tool List for ISP Product line | DHF379868 | LH |
| REF – 13 | ISP13 3rd Party items assessment checklist | DHF380838 | E |
| REF – 14 | ISP13 Compatibility Impact Analysis | D000819409 | E |
| REF – 15 | ISP13 URS-PRS Traceability | D001105259 | C |
| REF – 16 | ISP13 Cardiac suite DRS | D000856146 | I |
| REF – 17 | ISP13_TestSpecMR Cardiac Functional LV RV | D000996525 | H |
| REF – 18 | ISP13 Temporal Enhancement Verification Specification | D001104277 | E |
| REF – 19 | ISP13_CardiacMapping_Test Specification | D001104051 | D |
| REF – 20 | ISP13_TestSpecMRCardiacQflow | D000996541 | E |
| REF – 21 | ISP13_TestSpecMR Cardiac Functional Long Axis | D000996542 | E |
| REF – 22 | ISP13_MR Cardiac Suite Sanity_Unit Test Specification | D001001801 | K |
| REF – 23 | ISP13 MR Cardiac Suite-Spatial Enhancement Verification Spec | D000996623 | E |
| REF – 24 | ISP13 MR Cardiac suite Viewer Verification Spec | D000994617 | F |
| REF – 25 | ISP13 RMA MR Cardiac Suite | D000991970 | F |
| REF – 26 | ISP13 Risk Management Matrix Cardiac Suite | D000879224 | F |
| REF – 27 | ISP13_MR Cardiac Suite Traceability Matrix | D000996539 | H |
| REF – 28 | ISP13_Risk Management Analysis_MRApps & MMV | D001065109 | A |
| REF – 29 | ISP13_RMM_MRAPPS & MMV | D001065108 | A |
| REF – 30 | Test Specification and Report HFAAPPS | D000995159 | H |
| REF – 31 | Traceability Matrix CT Apps | D000993523 | H |
| REF – 32 | ISP13 RMA for CTapps | D000938106 | G |
| REF – 33 | ISP13 RMM General CTApps | D000892550 | G |
| REF – 34 | Test Specification and Report SystemHFA | D000996325 | D |
| REF – 35 | Traceability Matrix System HFA | D001159554 | C |
| REF – 36 | ISP13 CCA DRS | D000872156 | E |
| REF – 37 | ISP13 TAVI DRS | D000865239 | B |
| REF – 38 | ISP13_Verification Test Specification and Report CT Cardiac | D000946877 | J |
| REF – 39 | ISP13_Sanity Test Specification and Report CT Cardiac | D000998132 | A |
| REF – 40 | ISP13 Traceability Matrix CT Cardiac | D000946874 | H |
| REF – 41 | ISP13_Risk Management Matrix_CT Cardiac | D000946860 | B |
| REF – 42 | ISP13 RMA CT Cardiac | D000946872 | D |
| REF – 43 | ISP13 SMGOP DRS | D000882706 | F |
| REF – 44 | ISP13 DRS Spectral CT Viewer | D001085961 | B |
| REF – 45 | ISP13 Spectral General Viewing DRS | D000822200 | F |
| REF – 46 | AVW15 DRS Spectral CTMA | D001142845 | B |
| REF – 47 | Test Specification-Spectral CT Multiphase Analysis | D001093929 | F |
| REF – 48 | ISP13 Test Spec CT Spectral sCTV | D001092392 | K |
| REF – 49 | ISP13 Test spec CTSpectral sMG | D001046073 | I |
| REF – 50 | ISP13 Test Script Validation Report-CTSpectral Automation | D001047485 | E |
| REF – 51 | ISP13-Code TRR CTSpectral Automation | D001047484 | E |
| REF – 52 | ISP13 Study approved CTSpectral apps | D001104395 | B |
| REF – 53 | ISP13 Traceability Matrix for CTSpectral Applications | D001097887 | E |
| REF – 54 | ISP13 RMA CTSpectral Apps | D000952406 | E |
| REF – 55 | ISP13 RMM CTSpectral Apps | D000996304 | E |
| REF – 56 | ISP13 CTV Functional Regression Test Specification | D000999069 | D |
| REF – 57 | ISP13 NM Apps_Sanity Test Specification | D001059465 | C |
| REF – 58 | ISP13 NM Apps RMF Test Specification TRR | D001107914 | B |
| REF – 59 | AVW15_DRS_System_Installation | D001149811 | B |
| REF – 60 | ISP13_PVT_Installation & Upgrades_TestSpec | D001115079 | F |
| REF – 61 | ISP13 PVT Owned CT Apps sanity Test Spec | D001048147 | A |
| REF – 62 | ISP13_PVT_TestSpec | D001058943 | C |
| REF – 63 | PVT_AVW15_Installation and Upgrades_Traceability | D001157062 | B |
| REF – 64 | DRS for Remote Tools Suite Installation & Features Requirements for AVW / Console / Enterprise | D001137317 | A |
| REF – 65 | ISP13_TestSpecification_SystemFeatures-BLR | D000999142 | F |
| REF – 66 | AVW15_TestSpecificion_Serviceability | D001164936 | B |
| REF – 67 | AVW15_Traceability_SystemFeatures-BLR & Serviceability | D001164937 | C |
| REF – 68 | ISP13 System RMM | D000996286 | D |
| REF – 69 | ISP13 DRS System Performance | D001157240 | B |
| REF – 70 | AVW15 NFT Performance Test Specification | D001163851 | B |
| REF – 71 | AVW15_Consolidated Code TRR for NFTPerformanceAutomation | D001163854 | B |
| REF – 72 | AVW15_Automation Test Script Validation Report_NFT_Performance | D001163852 | B |
| REF – 73 | AVW15 Traceability Matrix_NFT | D001163855 | B |
| REF – 74 | ISP13 AVA DRS | D001083895 | C |
| REF – 75 | AVW15 DRS CT APPS BP | D001188993 | A |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

| Reference | Document Title | Document ID | Revision |
| --- | --- | --- | --- |
| REF – 1 | Aspect DRS | D001001539 | B |
| REF – 2 | CT Aspect RMM | D000892550 | G |
| REF – 3 | Test Specification_Aspect | D001157474 | A |
| REF – 4 | Design Verification Plan_Aspect | D001157489 | A |
| REF – 5 | Traceability Matrix CT ASPECTS | D001157485 | A |

#### RC Release

| Reference | Document Title | Document ID | Revision |
| --- | --- | --- | --- |
| REF – 1 | Aspect DRS | D001001539 | B |
| REF – 2 | CT ASPECTS RMM | D001157471 | B |
| REF – 3 | Test Specification_Aspect | D001157474 | A |
| REF – 4 | Design Verification Plan_Aspect | D001157489 | A |
| REF – 5 | Traceability Matrix CT ASPECTS | D001157485 | A |

### NM Suite

#### RFD Release

| Reference | Document Title | Document ID | Revision |
| --- | --- | --- | --- |
| REF – 1 | NM Apps Sanity_UnitTest Specification | D001033746 | A |
| REF – 2 | NM Apps Regression Test Specification TRR | D001033747 | B |
| REF – 3 | NM Apps_Traceability Matrix | D001033748 | B |
| REF – 4 | RMA_NM Apps | D001024262 | A |
| REF – 5 | NM Apps_RMM | D001024288 | A |

#### RC Release

| Reference | Document Title | Document ID | Revision |
| --- | --- | --- | --- |
| REF – 1 | NM Apps Sanity_UnitTest Specification | D001033746 | A |
| REF – 2 | NM Apps Regression Test Specification TRR | D001033747 | B |
| REF – 3 | NM Apps_Traceability Matrix | D001033748 | B |
| REF – 4 | RMA_NM Apps | D001024262 | A |
| REF – 5 | NM Apps_RMM | D001024288 | A |

## Abbreviations and Definitions

| Term | Definition |
| --- | --- |
| AVW | Advanced Visualization Workspace |
| BLR | Bengaluru |
| CAD RADS | Coronary Artery Disease Reporting and Data System |
| CT | Computed Tomography |
| DHF | Design History File |
| DRS | Detailed Requirement Specification |
| EV | End Verification |
| F2A | “FixToApprove” Defects in Resolved state. |
| HFA | Haifa |
| Invalid Defects | Defects closed with resolution – Work As Specified / Invalid or Uncontrolled Configuration / Does Not Apply / Not Enough Information / Obsolete Function / Rejected found during Verification Phase. |
| ISP | IntelliSpace Portal |
| MR | Magnetic Resonance |
| PIAB | Portal In A BOX |
| PRC | Product Release Committed |
| PRS | Product Requirements Specifications |
| PSO | Product Security Officer |
| Q&R | Quality & Regulatory |
| RFC | Request for Change |
| RFD | Release for Delivery |
| RMA | Risk Management Analysis |
| RMF | Risk Management File |
| RMM | Risk Management Matrix |
| SVer | Start Verification |
| TFS | Microsoft Visual Studio Team Foundation Server |
| VAL | Validation |

# Strategy

- Manual / Automation test execution of New Features

- Manual / Automation test execution of Enhanced Feature

- Manual / Automation test execution of Legacy Impacted Features

- Safety requirements that provide mitigations for known risks

- Expert tests that will be done by Application’s specialists.

# Test Design

## Test Protocol

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

| Test Protocol | Document ID |
| --- | --- |
| ISP13 MR Cardiac suite Viewer Verification Spec | D000994617 |
| ISP13 MR Cardiac suite Functional LVRV Verification Spec | D000996525 |
| ISP13 MR Cardiac suite QFlow Verification Spec | D000996541 |
| ISP13 MR Cardiac suite Functional Long Axis Verification Spec | D000996542 |
| ISP13 Test Specification HFA (CT Liver) | D000995159 |
| ISP13 Test Specifications CT Cardiac | D000946877 |
| ISP13 Sanity Spec CT Cardiac | D000998132 |

RMM Tests Execution in ISP13 VAL-3 Development Phase:

| Test Protocol | Document ID |
| --- | --- |
| ISP13_PVT_TestSpec | D001058943 |
| ISP13_NM Apps_Sanity Test Specification | D001059465 |
| ISP13_NM Apps_Regression Test Specification | D001059467 |
| ISP13 Test Specification HFA | D000995159 |
| ISP13 Test Specifications CT Cardiac | D000946877 |
| ISP13 Test spec CTSpectral sMG | D001046073 |
| ISP13 CTV Functional Regression Test Specification | D000999069 |

#### ISP13 VAL-3 Release

| Test Protocol | Document ID |
| --- | --- |
| ISP13_TestSpecMR Cardiac Functional LV RV | D000996525 |
| ISP13 Temporal Enhancement Verification Specification | D001104277 |
| ISP13 MR Cardiac Suite Mapping Functionality Test specification | D001104051 |
| ISP13 LoBI Regression Verification Spec | D001104696 |
| ISP13 NM Apps_Sanity Test Specification | D001059465 |
| ISP13 NM Apps RMF Test Specification TRR | D001107914 |
| ISP13_Verification Test Specification and Report CT Cardiac | D000946877 |
| Test Specification and Report HFAAPPS | D000995159 |
| ISP13 Test Specification-Spectral CT Multiphase Analysis | D001093929 |
| ISP13 Test Spec CT Spectral sCTV | D001092392 |
| ISP13 Test spec CTSpectral sMG | D001046073 |
| ISP13 Test Script Validation Report-CTSpectral Automation | D001047485 |
| ISP13_PVT_TestSpec | D001058943 |
| ISP13_TestSpecification_SystemFeatures-BLR | D000999142 |

RMM Tests Execution in AVW15 RFD Development Phase:

| Test Protocol | Document ID |
| --- | --- |
| ISP13 MRCardiac suite Functional LVRV Verification spec | D000996525 |
| ISP13 MR Cardiac Suite-QFlow Verification Specification | D000996541 |
| ISP13 MR Cardiac Suite-Functional Long Axis Verification Specification | D000996542 |
| ISP13 MR Cardiac Suite Unit/Sanity Test Specification | D001001801 |
| ISP13_CardiacMapping_Test Specification | D001104051 |
| ISP13 Temporal Enhancement Test Specification and Report | D001104277 |
| ISP13 MR Cardiac Suite-Spatial Enhancement Verification Spec | D000996623 |
| ISP13 MR Cardiac Hazard Verification Spec Legacy Application | D001064798 |
| ISP13 MR Cardiac suite Viewer Verification Spec | D000994617 |
| ISP13 Regression Spec LOBI | D001104696 |
| Test Specification HFA | D000995159 |
| ISP13 Automation Validation Report for CT Apps and BCKHFA | D000996313 |
| ISP13 Sanity Spec CT Cardiac | D000998132 |
| ISP13 Test Specifications CT Cardiac | D000946877 |
| ISP13 Study approved CTSpectral apps | D001104395 |
| ISP13 Test Script Validation Report-CTSpectral Automation | D001047485 |
| Test Specification-Spectral CT Multiphase Analysis | D001093929 |
| ISP13 Test Spec CT Spectral sCTV | D001092392 |
| ISP13 Test spec CTSpectral sMG | D001046073 |
| ISP13 CTV Functional Regression Test Specification | D000999069 |
| ISP13 NM Apps_Sanity Test Specification | D001059465 |
| ISP13 NM Apps RMF Test Specification TRR | D001107914 |
| ISP13 PVT Owned CT Apps sanity Test Spec | D001048147 |
| ISP13_PVT_TestSpec | D001058943 |
| ISP13_TestSpecification_SystemFeatures-BLR | D000999142 |

#### AVW15 RFD Release

| Test Protocol | Document ID |
| --- | --- |
| ISP13_TestSpecMR Cardiac Functional LV RV | D000996525 |
| ISP13 Temporal Enhancement Verification Specification | D001104277 |
| ISP13_CardiacMapping_Test Specification | D001104051 |
| ISP13_TestSpecMRCardiacQflow | D000996541 |
| ISP13_TestSpecMR Cardiac Functional Long Axis | D000996542 |
| ISP13_MR Cardiac Suite Sanity_Unit Test Specification | D001001801 |
| ISP13 MR Cardiac Suite-Spatial Enhancement Verification Spec | D000996623 |
| ISP13 MR Cardiac suite Viewer Verification Spec | D000994617 |
| Test Specification and Report HFAAPPS | D000995159 |
| Test Specification and Report SystemHFA | D000996325 |
| ISP13_Verification Test Specification and Report CT Cardiac | D000946877 |
| ISP13_Sanity Test Specification and Report CT Cardiac | D000998132 |
| Test Specification-Spectral CT Multiphase Analysis | D001093929 |
| ISP13 Test Spec CT Spectral sCTV | D001092392 |
| ISP13 Test spec CTSpectral sMG | D001046073 |
| ISP13 CTV Functional Regression Test Specification | D000999069 |
| ISP13 NM Apps_Sanity Test Specification | D001059465 |
| ISP13 NM Apps RMF Test Specification TRR | D001107914 |
| ISP13_PVT_Installation & Upgrades_TestSpec | D001115079 |
| ISP13 PVT Owned CT Apps sanity Test Spec | D001048147 |
| ISP13_PVT_TestSpec | D001058943 |
| ISP13_TestSpecification_SystemFeatures-BLR | D000999142 |
| AVW15_TestSpecificion_Serviceability | D001164936 |
| AVW15 NFT Performance Test Specification | D001163851 |

#### AVW15 RC Release

| Test Protocol | Document ID |
| --- | --- |
| ISP13_TestSpecMR Cardiac Functional LV RV | D000996525 |
| ISP13 Temporal Enhancement Verification Specification | D001104277 |
| ISP13_CardiacMapping_Test Specification | D001104051 |
| ISP13_TestSpecMRCardiacQflow | D000996541 |
| ISP13_TestSpecMR Cardiac Functional Long Axis | D000996542 |
| ISP13_MR Cardiac Suite Sanity_Unit Test Specification | D001001801 |
| ISP13 MR Cardiac Suite-Spatial Enhancement Verification Spec | D000996623 |
| ISP13 MR Cardiac suite Viewer Verification Spec | D000994617 |
| Test Specification and Report HFAAPPS | D000995159 |
| Test Specification and Report SystemHFA | D000996325 |
| ISP13_Verification Test Specification and Report CT Cardiac | D000946877 |
| ISP13_Sanity Test Specification and Report CT Cardiac | D000998132 |
| Test Specification-Spectral CT Multiphase Analysis | D001093929 |
| ISP13 Test Spec CT Spectral sCTV | D001092392 |
| ISP13 Test spec CTSpectral sMG | D001046073 |
| ISP13 CTV Functional Regression Test Specification | D000999069 |
| ISP13 NM Apps_Sanity Test Specification | D001059465 |
| ISP13 NM Apps RMF Test Specification TRR | D001107914 |
| ISP13_PVT_Installation & Upgrades_TestSpec | D001115079 |
| ISP13 PVT Owned CT Apps sanity Test Spec | D001048147 |
| ISP13_PVT_TestSpec | D001058943 |
| ISP13_TestSpecification_SystemFeatures-BLR | D000999142 |
| AVW15_TestSpecificion_Serviceability | D001164936 |
| AVW15 NFT Performance Test Specification | D001163851 |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

| Test Protocol | Document ID |
| --- | --- |
| Test Specification_Aspect | D001157474 |

#### RC Release

| Test Protocol | Document ID |
| --- | --- |
| Test Specification_Aspect | D001157474 |

### NM Suite

#### RFD Release

| Test Protocol | Document ID |
| --- | --- |
| ISP12 NM Apps Sanity_UnitTest Specification | D001033746 |
| ISP12.4 NM Apps Regression Test Specification TRR | D001033747 |

#### RC Release

| Test Protocol | Document ID |
| --- | --- |
| ISP12 NM Apps Sanity_UnitTest Specification | D001033746 |
| ISP12.4 NM Apps Regression Test Specification TRR | D001033747 |

## Verification Activities Approach Overview

The Verification Plan gives the overview of the complete tests planned for:

Scope:

- Sanity Tests

- Execution of the functional tests for New and Enhanced Features.

- Execution of the functional tests for Legacy Impacted features – Regression Tests

- Execution of RMF test.

- Impacted feature test for Defect Verification.

- Expert Testing

- Anti-Virus Scan.

Verification Types

| Types of Verification | When is it done? | How is it documented? |
| --- | --- | --- |
| Unit Verification | Prior to SVER Phase | In Unit Test report |
| Static Verification (Test Spec) | Prior to SVER Phase | In Test specifications document. |
| Code Reviews | Prior to SVER Phase | In Code Technical Review Record |
| Integration Testing | During Verification | In Verification report Note: Integration Testing is done by executing the Sanity tests on SVer & Final ISO for Verification. This is the integration automation sanity testing report finalized by the PVT team. |
| System Testing | During Verification | In Verification Report |
| Expert Testing | During Verification | In Expert Test Report |
| Anti-Virus Scan | During Verification | In Anti-Virus Scan Report |

## Methods Overview

| Test Type | Purpose of Test |
| --- | --- |
| Full functionality test | A comprehensive list of test cases per feature which is divided to many short test cases. Each test case covers a specific function/scenario in the feature. The full functionality test (which includes all of the feature's functionality test cases) covers all of the detailed requirements and (in case all test cases are passed) provides us with the assurance that this feature is working as required. |
| RMF tests | RMF (Risk management file) tests are tests that are specifically designed to check scenarios that may lead to hazards. The tests are designed according to the RMF file, and in general check calculations, measurements, patient details, information, and other scenarios that might lead to misdiagnosis (in case of SW defects). |
| Regression test | Regression test is a functionality test of a specific part of the application, which might be impacted due to new feature implementation / Bug fixes / RFCs. Regression test is derived from the functionality test set, and usually includes a part of the full functionality test cases. |
| Sanity test | Sanity test is a wide but shallow test that gives an overview of the build's maturity. It contains simple, random operations for each feature, and (in case everything passes) indicates that the build is mature enough for functionality and deeper testing. |
| Freehand test / Exploratory test | A test which is not written with formal, technical steps, but is a dynamic test on which the tester decided what to test according to the feature reactions (for example, if the tester finds a serious defect on a specific part of the feature, he/she might decide to test some more scenarios that might be impacted by the problematic behavior and might cause additional defects). |
| Expert test | Expert testing is done by application specialist/ system engineer/ system architect to test the Clinical/Application workflow as used by end users. The testing will be done by experts based on the knowledge of the expert and does not need any preapproved or predefined test protocol. Expert test report will capture the feedback. |
| Anti-Virus Scan | Any release/package to the field should pass Anti-Virus scan and the results of each release/package need to be documented in the verification report (ICAP-T-030010.10 System Malware Scan Verification Report). |
| Automation test | A test running automatically by a software tool (Test complete) that controls the execution of tests and the comparison of actual outcomes with predicted outcomes. The outcome results are reviewed and approved by a manual tester. |

### Coverage Strategy

Plan test scenarios based on the dataset variance. Also ensure that the following steps are followed
1) Create System workflow scenarios
2) Enterprise level scenarios to be documented
3) Start collating Operational Profiling of Portal System 
4) Perform Workflow testing (majorly used use cases for enterprise and applications).
5) Perform a test of the final installation media over each type of deliverable HW for the version:
        a) OS installation media shall be tested over each of the deliverable HW for the version
        b) 1-Click installation on the final media (Portal / Concerto) shall be tested once over relevant HW        
              deliverable for the version.
6) Execute above tests

#### Advance Visualization Workspace 15

##### ISP13 VAL-2 Release: Detailed Verification Coverage

| Features / Applications | Test Owner | Verification Method | Verification Strategy |
| --- | --- | --- | --- |
| MR Cardiac Suite – Viewer | BLR Apps – MR Apps | Full Functionality + RMM Tests | Viewer of MR Cardiac Suite is enhanced with new features hence full functionality and RMM test is planned. |
| MR Cardiac Suite – Functional LV & RV | BLR Apps – MR Apps | Full Functionality + RMM Tests | Functional LV & RV application of MR Cardiac Suite is enhanced with new features hence full functionality and RMM test is planned. |
| MR Cardiac Suite – Qflow | BLR Apps – MR Apps | Full Functionality + RMM Tests | QFlow application of MR Cardiac Suite is enhanced with new features hence full functionality and RMM test is planned. |
| MR Cardiac Suite – Spatial Enhancement | BLR Apps – MR Apps | Full Functionality + RMM Tests | Spatial Enhancement application of MR Cardiac Suite is enhanced with new features hence full functionality and RMM test is planned. |
| MR Cardiac Suite – Functional LA | BLR Apps – MR Apps | Full Functionality + RMM Tests | Functional LA application of MR Cardiac Suite is enhanced with new features hence full functionality and RMM test is planned. |
| CT Liver | HFA Apps | Full Functionality | CT Liver application of CT Modality is enhanced with new features hence full functionality test is planned. |
| CCA | BLR Apps – Shina Apps | Full Functionality + RMM Tests | CCA application of CT Modality is enhanced with a new feature hence full functionality and RMM test is planned. |

RMM Tests Execution in ISP13 VAL-3 Development Phase:

| Features / Applications | Test Owner | Verification Method | Verification Strategy |
| --- | --- | --- | --- |
| RMM | - MR Apps - NM Apps - CT Apps - CT Cardiac Apps - CT Spectral Apps - CT Viewing | RMF Tests | As “RMF Scenarios” are critical functionality, RMF test is planned. |

##### ISP13 VAL-3 Release: Detailed Verification Coverage

| Features / Applications | Test Owner | Verification Method | Verification Strategy |
| --- | --- | --- | --- |
| MR Cardiac Suite – Temporal Analysis | BLR Apps – MR Apps | Full Functionality + RMM Tests | Temporal Analysis application of MR Cardiac Suite is enhanced with new features hence full functionality and RMM test is planned. |
| MR Cardiac Suite – Functional LV & RV | BLR Apps – MR Apps | Full Functionality + RMM Tests | Functional LV & RV application of MR Cardiac Suite is enhanced with new features hence full functionality and RMM test is planned. |
| MR Cardiac Suite – Mapping | BLR Apps – MR Apps | Full Functionality + RMM Tests | Mapping application of MR Cardiac Suite is enhanced with new features hence full functionality and RMM test is planned. |
| LOBI CVS | BLR Apps – MR Apps | Full Functionality + RMM Tests | LOBI application is enhanced with new feature hence full functionality of new feature and RMM test is planned. |
| MR Cardiac Suite | BLR Apps – MR Apps | RMM + Regression | MR Cardiac Suite application specific RMM & Regression tests are planned. |
| MR Apps | BLR Apps – MR Apps | RMM + Regression | MR applications specific RMM & Regression tests are planned. |
| NM Apps | BLR Apps – NM Apps | RMM + Sanity | NM applications specific RMM & Sanity tests are planned. |
| CAD RADS | BLR Apps – Shina Apps | Regression | CAD RADS regression tests are planned. |
| CCA & TAVI Legacy | BLR Apps – Shina Apps | RMM | CCA & TAVI Legacy RMM tests are planned. |
| AutoCS | CT Apps | Full Functionality + RMM + Regression | AutoCS application of CT Modality is enhanced with new features hence full functionality of new features, regression and RMM tests are planned. |
| sAVA | CT Apps | Full Functionality + RMM + Regression | sAVA application of CT Modality is enhanced with new features hence full functionality of new features, regression and RMM tests are planned. |
| AVA regression | CT Apps | Regression | AVA regression of Legacy area from previous versions |
| Trauma Legacy RMF | CT Apps | RMM | RMF tests for legacy Trauma application is planned. |
| sMGOP/LMG 3D and MPR | CT Spectral Apps | Full Functionality + RMM | sMG/LMG enhanced with new features hence full functionality and RMM tests are planned. |
| sCTV RFCs | CT Spectral Apps | Full Functionality + RMM | New RFCs introduced, hence full functionality tests and RMM tests are planned. |
| Legacy CT Apps | PVT | RMM | RMF tests for legacy CT Applications are planned. |
| Legacy MR Apps | PVT | RMM | RMF tests for legacy MR Applications are planned. |
| Legacy MR Cardiac Apps | PVT | RMM | RMF tests for legacy MR Cardiac Applications are planned. |
| Legacy NM Apps | PVT | RMM | RMF tests for legacy NM Applications are planned. |
| NLog | System BLR | Full Functionality | NLog is a new feature hence full functionality tests are planned. |
| System Features | System BLR | Regression | Regression tests for System features are planned. |
| System RMM | System BLR | Regression | RMF tests for System features are planned. |
| PMCF for MR cardiac | System HFA | Full Functionality | Full functionality tests of PMCF for MR Cardiac Suite, relevant to VAL3 are planned. |
| Labeling | System HFA | Regression | Regression tests for Labeling feature, relevant to VAL3 are planned. |
| Installation | PVT | Regression | Regression tests for Installation testing are planned. |

RMM Tests Execution in AVW15 RFD Development Phase:

| Features / Applications | Test Owner | Verification Method | Verification Strategy |
| --- | --- | --- | --- |
| RMM | - MR Apps - CT Apps - CT Cardiac Apps - CT Spectral Apps - CT Viewing - NM Apps - CT, MR, NM Legacy Apps - System Features | RMF Tests | As “RMF Scenarios” are critical functionality, RMF test is planned. |

##### AVW15 RFD Release: Detailed Verification Coverage

| Features / Applications | Test Owner | Verification Method | Verification Strategy |
| --- | --- | --- | --- |
| Findings, Bookmarks & Save Functionality | MR Cardiac Suite | Full Functionality + Regression | Findings, Bookmarks & Save Functionality for MR Cardiac Suite applications Full functionality & Regression tests are planned. |
| RMM Tests | MR Cardiac Suite | RMF | Safety tests of MR Cardiac Suite applications are planned. |
| Regression Tests | MR Cardiac Suite | Regression | Regression tests for impacted MR Cardiac Suite applications are planned. |
| Performance Tests | MR Cardiac Suite | Regression | Performance tests for MR Cardiac Suite applications are planned. |
| RMM Tests | MR Applications | RMF | Safety tests of MR Cardiac Suite applications are planned. |
| Standard Reporting | CT Applications | Full Functionality | Standard Reporting application of CT Modality is enhanced with new features hence full functionality of new features planned be executed on LNA, CS and AVA applications |
| CS Legacy RMF | CT Applications | RMF | RMF tests for legacy Trauma application is planned. |
| BP Legacy RMF | CT Applications | RMF | RMF tests for legacy Trauma application is planned. |
| 3D Modeling Legacy RMF | CT Applications | RMF | RMF tests for legacy Trauma application is planned. |
| BMD Legacy RMF | CT Applications | RMF | RMF tests for legacy Trauma application is planned. |
| VC Legacy RMF | CT Applications | RMF | RMF tests for legacy Trauma application is planned. |
| LNA Legacy RMF | CT Applications | RMF | RMF tests for legacy Trauma application is planned. |
| Regression Tests | CT Applications | Regression | Full coverage of PRS |
| AVW Labeling | System Features HFA | Full Functionality | New Functionality + Regression tests for Labeling feature. |
| ZFP | System Features HFA | Full Functionality | New Functionality + Regression tests for ZFP feature, |
| Application name rebranding | System Features HFA | Full Functionality | New Functionality + Regression tests for Rebranding feature, |
| 3rd party application in Client | System Features HFA | Full Functionality | New Functionality + Regression tests for Client installation feature, |
| Regression Tests | System Features HFA | Regression | Full coverage of PRS |
| CAD RADS & Findings Management | CT Cardiac Applications | Full Functionality + Regression | CAD RADS & Findings Management for CT Cardiac applications Full functionality & Regression tests are planned. |
| Discrete Report – CCA and TAVI | CT Cardiac Applications | Full Functionality | Discrete Report for CCA and TAVI applications Full functionality tests are planned. |
| RMM Tests | CT Cardiac Applications | RMF | Safety tests of CT Cardiac Applications are planned. |
| sMGOP/LMG 3D and 2D support | CT Spectral Applications | Full Functionality + RMF | sMG/LMG enhanced with new features hence full functionality and RMM tests are planned. |
| sCTV RFCs | CT Spectral Applications | Full Functionality + RMF | New RFCs introduced, hence full functionality tests and RMM tests are planned. |
| RMM Tests | CT Spectral Applications | RMF | Safety tests of CT Spectral Applications are planned. |
| CT Viewing RMM | CT Viewing | RMF | Safety tests of CTViewing are planned. |
| NM Applications 3rd Party : 1. AutoQuant 2. Corridor4DM 3. ECTB 4. NeuroQ 5. Mirada XD | NM Applications | Regression | Regression tests for 3rd party NM Applications are planned. |
| Legacy Applications -CT | CT, MR, NM Legacy Applications | Sanity, Regression & RMM Tests | Sanity, Regression & RMM tests for legacy CT Applications are planned. |
| Legacy Applications -NM | CT, MR, NM Legacy Applications | Sanity, Regression & RMM Tests | Sanity, Regression & RMM tests for legacy NM Applications are planned. |
| Legacy Applications -MR | CT, MR, NM Legacy Applications | Sanity, Regression & RMM Tests | Sanity, Regression & RMM tests for legacy MR Applications are planned. |
| Legacy Applications -MRC | CT, MR, NM Legacy Applications | Sanity, Regression & RMM Tests | Sanity, Regression & RMM tests for legacy MRC Applications are planned. |
| 3rd Party Decoupling | Installation, Upgrades & Security | Full functionality + regression | 3rd Party Decoupling Full functionality & Regression tests are planned. |
| Client Installer Re-design | Installation, Upgrades & Security | Full functionality + regression | Client Installer Re-design Full functionality & Regression tests are planned. |
| Installation, upgrade & Security | Installation, Upgrades & Security | Regression tests | Regression tests for Installation, Upgrades & Security tests are planned. |
| DynaCad 5.0 Full functionality | System Features-BLR | Full functionality + regression | DynaCad 5.0 Full functionality & Regression tests are planned. |
| License – KGS & EKGS | System Features-BLR | Regression | Regression tests for Licensing feature are planned. |
| System Features | System Features-BLR | Regression | Regression tests for impacted System Features are planned |
| RMM Tests | System Features-BLR | RMM | Safety tests of System Features are planned. |
| Localization | System Features-BLR | Regression | Localization tests System Features are Planned. |
| Compatibility Tests | System Features-BLR | Regression | Compatibility Tests for supporting reference system are planned. |
| IOT | Serviceability | Full functionality | As IOT is a new feature, full functionality tests are planned. |
| Serviceability Features | Serviceability | Regression | Regression tests for impacted Serviceability Features are planned. |
| Performance | Non-Functional Tests | Full Functionality | Manual & Automation tests execution of applications performance are planned. |

##### AVW15 RC Release: Detailed Verification Coverage

| Features / Applications | Test Owner | Verification Method | Verification Strategy |
| --- | --- | --- | --- |
| MR Cardiac suite Functional LVRV | MR Cardiac Suite | RMM and Regression | Regression Test, Sanity and RMM Test case execution |
| MR Cardiac suite Qflow | MR Cardiac Suite | RMM and Regression | Regression Test, Sanity and RMM Test case execution |
| MR Cardiac suite Functional LALV | MR Cardiac Suite | RMM and Regression | Regression Test, Sanity and RMM Test case execution |
| MR Cardiac suite Temporal Enhancement | MR Cardiac Suite | RMM and Regression | Regression Test, Sanity and RMM Test case execution |
| MR Cardiac suite Spatial Enhancement | MR Cardiac Suite | RMM and Regression | Regression Test, Sanity and RMM Test case execution |
| MR Cardiac suite Viewer | MR Cardiac Suite | RMM and Regression | Regression Test, Sanity and RMM Test case execution |
| Lobi and Liver | MR Cardiac Suite | Regression | Regression Test case execution |
| MR Apps | MR Apps | Regression | Regression Test case execution |
| Standard Reporting | CT Applications | Regression | Regression tests for impacted areas |
| QR RMF | CT Applications | RMF | RMF tests for legacy QR application is planned. |
| CT Pulmo RMF | CT Applications | RMF | RMF tests for legacy CT Pulmo application is planned. |
| IFU RMF | CT Applications | RMF | RMF tests for legacy IFU application is planned. |
| BP | CT Applications | Updated functionality + Regression | Regression tests for impacted areas |
| CT Applications | CT Applications | Regression | Regression tests for impacted areas |
| System HFA | System HFA | Regression | Regression tests for impacted areas |
| Labeling | System HFA | Regression | Regression tests for impacted areas |
| CAD-RADS and Findings Management | CT Cardiac | Regression, functional and RMF | Regression test (Includes defect Verification) Functional test, RMF execution, Sanity and Upgrades |
| CCA RFCs | CT Cardiac | Regression, functional and RMF | Regression test (Includes defect Verification) Functional test, RMF execution, Sanity and Upgrades |
| sCTV and SMGOP new features | CT Spectral Applications | Functional and Regression | Functional Regression tests for impacted areas |
| CT Viewing - Angio CT | CT Viewing | Regression | Regression of CT Viewing is planned. |
| NM Applications 3rd Party : 1. AutoQuant 2. Corridor4DM 3. ECTB 4. NeuroQ 5. Mirada XD | NM Applications | Regression | Regression tests for 3rd party NM Applications are planned. |
| KnowledgeScape | NM Applications | Sanity and Regression | Sanity and Regression on the new content for RC |
| Legacy Applications -CT | CT, MR, NM Legacy Applications | Sanity, Regression Tests | Sanity, Regression tests for legacy CT Applications are planned. |
| Legacy Applications -NM | CT, MR, NM Legacy Applications | Sanity, Regression Tests | Sanity, Regression tests for legacy NM Applications are planned. |
| Legacy Applications -MR | CT, MR, NM Legacy Applications | Sanity, Regression & RMM Tests | Sanity, Regression & RMM tests for legacy MR Applications are planned. |
| Legacy Applications -MRC | CT, MR, NM Legacy Applications | Sanity, Regression Tests | Sanity, Regression tests for legacy MRC Applications are planned. |
| Installation, upgrade & Security | Installation, Upgrades & Security | Regression tests | Regression tests for Installation, Upgrades & Security tests are planned. |
| DynaCad 5.1 | System Features-BLR | regression | DynaCad 5.0 Regression tests are planned. |
| License – KGS & EKGS | System Features-BLR | Regression | Regression tests for Licensing feature are planned. |
| System Features | System Features-BLR | Regression | Regression tests for impacted System Features are planned |
| PMT Client Auto upgrade | System Features-BLR | Full functionality | Full functionality tests for PMT Client Auto upgrade Feature are Planned. |
| Compatibility Tests | System Features-BLR | Regression | Compatibility Tests for supporting reference system are planned. |
| Axeda Split package and SPLUNK | Serviceability | Regression | Regression tests are planned. |
| IOT | Serviceability | Full functionality | Full functionality tests for IOT Feature are Planned. |
| Serviceability Features | Serviceability | Regression | Regression tests for impacted Serviceability Features are planned. |
| Performance | Non-Functional Tests | Full Functionality | Manual & Automation tests execution of applications performance are planned. |

#### CT ASPECT Scoring (CT ASPECTS)

##### RFD Release

| Features / Applications | Test Owner | Verification Method | Verification Strategy |
| --- | --- | --- | --- |
| CT Aspects | CT Applications | Regression | Regression tests for CT ASPECTS application are planned. |

##### RC Release

| Features / Applications | Test Owner | Verification Method | Verification Strategy |
| --- | --- | --- | --- |
| CT Aspects | CT Applications | Sanity | Regression tests for CT ASPECTS application are planned in Automation. |

#### NM Suite

##### RFD Release

| Features / Applications | Test Owner | Verification Method | Verification Strategy |
| --- | --- | --- | --- |
| NM Planar Suite | NM Applications | Sanity, Regression & RMM Tests | Sanity, Regression & RMM tests for NM Planar Suite are planned. |
| NM JETPack | NM Applications | Sanity, Regression & RMM Tests | Sanity, Regression & RMM tests for NM JETPack are planned. |
| NM QA QC Tools | NM Applications | Sanity, Regression & RMM Tests | Sanity & Regression tests for NM QA QC Tools are planned. |

##### RC Release

| Features / Applications | Test Owner | Verification Method | Verification Strategy |
| --- | --- | --- | --- |
| NM Planar Suite | NM Applications | Sanity | Sanity for NM Planar Suite are planned. |
| NM JETPack | NM Applications | Sanity | Sanity for NM JETPack are planned. |
| NM QA QC Tools | NM Applications | Sanity | Sanity for NM QA QC Tools are planned. |

- All testing methods that have been chosen for each phase, have been selected carefully, taking into consideration all available information on the changes that perform within the applications from the software side, project management and application specialists.

- Test results are marked –

  - “PASSED” – when the executed scenario had no steps which specified as failed.

  - “FAILED” – when the executed scenario had steps which specified as failed.

  - “N/A” – when the planned test case is not executable.

- Test results are marked –

  - “PASSED” – when the actual result is matching expected result.

  - “FAILED” – when the actual result is not matching expected result.

  - “NONE” – when the step is N/A or all steps of a test case are not executable.

- Tests planned for Verification are required to use the datasets from one of the below mentioned dataset repository path:

  - BLR: \\ingbtcpic5vw113\ICAP_BLR Repository\ISP_Verification_Dataset

  - HFA: R:\Image Types\CT PHILIPS

  - BLR: \\ingbtcpic5vw113\ICAP_BLR Repository\ISP_Verification_Dataset

  - HFA: R:\Image Types\ICAP HFA Dataset Repository\ISP Standard Verification Datasets

Tests Execution is not location specific. Tests designed to use the dataset from BLR location can be executed in HFA location and vice versa. In such case the dataset mentioned in Description / Expected Result of the test case shall be copied to dataset repository of specific location and Actual results shall be updated accordingly.

### Requirements Traceability Strategy

#### Advance Visualization Workspace 15

Requirements Traceability matrix for below features / enhancement is documented in:

##### ISP13 VAL-2 Release:

| Scope | Traceability Document | DHF ID |
| --- | --- | --- |
| MR Cardiac Suite: - Viewer - Functional LV & RV - Qflow - Spatial Enhancement - Functional LA | ISP13 MR Cardiac Suite Traceability Matrix | D000996539 |
| CT Liver | ISP13 Traceability Matrix HFA Apps | D000993523 |
| CCA | ISP13 Traceability Matrix CT Cardiac | D000946874 |

##### ISP13 VAL-3 Release:

| Scope | Traceability Document | DHF ID |
| --- | --- | --- |
| MR Cardiac Suite: - Temporal Analysis - Functional LV RV - Mapping | ISP13 MR Cardiac Suite Traceability Matrix | D000996539 |
| LOBI CVS | ISP13 Traceability Matrix LoBI | D001105932 |
| sAVA | Traceability Matrix CT Apps | D000993523 |
| Auto CS | Traceability Matrix CT Apps | D000993523 |
| CT Spectral Apps | ISP13 Traceability Matrix for CTSpectral Applications | D001097887 |

##### AVW15 RFD Release:

| Scope | Traceability Document | DHF ID |
| --- | --- | --- |
| Findings, Bookmarks & Save Functionality for MR Cardiac Suite | ISP13_MR Cardiac Suite Traceability Matrix | D000996539 |
| Standard Reporting | Traceability Matrix CT Apps | D000993523 |
| ZFP | Traceability Matrix System HFA | D001159554 |
| 3rd party application | Traceability Matrix System HFA | D001159554 |
| Labeling and Rebranding | Traceability Matrix System HFA | D001159554 |
| CCA and TAVI Discrete Report | ISP13 Traceability Matrix CT Cardiac | D000946874 |
| CT Spectral Applications | ISP13 Traceability Matrix for CTSpectral Applications | D001097887 |
| Installations and Upgrade | PVT_AVW15_Installation and Upgrades_Traceability | D001157062 |
| IOT | AVW15_Traceability_SystemFeatures-BLR & Serviceability | D001164937 |
| Applications Performance | AVW15 Traceability Matrix_NFT | D001163855 |

##### AVW15 RC Release:

| Scope | Traceability Document | DHF ID |
| --- | --- | --- |
| Findings, Bookmarks & Save Functionality for MR Cardiac Suite | ISP13_MR Cardiac Suite Traceability Matrix | D000996539 |
| Standard Reporting | Traceability Matrix CT Apps | D000993523 |
| ZFP | Traceability Matrix System HFA | D001159554 |
| 3rd party application | Traceability Matrix System HFA | D001159554 |
| Labeling and Rebranding | Traceability Matrix System HFA | D001159554 |
| CCA and TAVI Discrete Report | ISP13 Traceability Matrix CT Cardiac | D000946874 |
| CT Spectral Applications | ISP13 Traceability Matrix for CTSpectral Applications | D001097887 |
| Installations and Upgrade | PVT_AVW15_Installation and Upgrades_Traceability | D001157062 |
| IOT | AVW15_Traceability_SystemFeatures-BLR & Serviceability | D001164937 |
| Applications Performance | AVW15 Traceability Matrix_NFT | D001163855 |

#### CT ASPECT Scoring (CT ASPECTS)

##### RFD Release:

| Scope | Traceability Document | DHF ID |
| --- | --- | --- |
| None |  |  |

##### RC Release

| Scope | Traceability Document | DHF ID |
| --- | --- | --- |
| None |  |  |

#### NM Suite

##### RFD Release:

| Scope | Traceability Document | DHF ID |
| --- | --- | --- |
| None |  |  |

##### RC Release:

| Scope | Traceability Document | DHF ID |
| --- | --- | --- |
| None |  |  |

## Tools Overview

The following is the list of tools that will be used during the project phase.

| Tool Name | Version | Purpose of Use |
| --- | --- | --- |
| TFS | Refer DHF379868 – ISP Tool List | Requirements, Test & Defect Management Tool. |
| PTC Windchill |  | Document Management Tool |
| Test Complete |  | Automation testing tool |
| Test Execute |  | Automation testing tool |
| Selenium |  | Automation testing tool |

Note: All the tools used in the activity of verification has been validated prior to usage for its intended usage.

## Test Environment Overview

### Advance Visualization Workspace 15

#### Target hardware

##### ISP13 VAL-2 Release:

- Demo Laptop – HP ZBook G3

- Server HW Configuration – Dell R640

- Workstation – HP Z4 (For CT Liver Application only)

##### ISP13 VAL-3 Release:

- Demo Laptop – HP ZBook G3

- Server HW Configuration – HP DL360

- Workstation – HP Z440

##### AVW15 RFD Release:

###### Execution Plan on Supported Hardware:

|  | Client Server |  |  | PIAB (IX / LX) |  |  | Demo Laptop | Conc erto | Virtualization |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Applications / Features | DELL T/R 640 | HP DL-360 Gen9 (Rack) | HP ML-350 Gen9 (Tower) | HP Z440 | HP Z4 (W-2145 Processor) | HP Z4 (W-2245 Processor) | HP ZBook G3 | VMWare virtualization solution | VMware Virtual Machine (ESXi) | Microsoft Hyper V | Nutanix |
| MR Cardiac Suite | X | X |  | X |  |  | X | X |  |  |  |
| MR Applications |  | X |  |  |  |  |  |  |  |  |  |
| CT Applications | X |  |  |  |  | X |  | X | X |  |  |
| System Features HFA | X |  |  |  |  | X |  | X | X | X |  |
| CT Cardiac Applications | X | X | X | X |  |  | X | X |  |  |  |
| CT Spectral Applications |  |  | X | X |  |  |  |  |  |  |  |
| CT Viewing |  | X |  |  |  |  |  |  |  |  |  |
| NM Applications |  | X |  | X |  |  |  |  |  |  |  |
| CT, MR, NM Legacy Applications | X | X |  |  | X |  |  |  |  |  |  |
| Installation, Upgrades & Security | X | X | X | X | X | X | X | X | X | X | X |
| System Features BLR | X | X |  | X |  |  |  | X |  |  |  |
| Serviceability | X | X |  | X |  |  |  | X |  |  |  |
| Non-Functional Tests |  | X |  |  |  |  |  |  |  |  |  |

###### Hardware Specification:

| Sl. No | Hardware Model | CPU (Model and number of CPUs) | RAM (Model & Size) | Graphic Card (Where available) | NIC – Model of Network Interface Card | Storage: Model, No. of drives & capacity |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Dell PowerEdge R640 | Model: Intel Xeon Gold 6136 3.0G, 12C/24T, 10.4GT/s 2UPI, 24.75M Cache, Turbo, HT (150W) DDR4-2666, OEM XL No. Of CPUs: 1 | Model: 16GB RDIMM, 2666MT/s, Dual Rank Size: 32 GB | Default on Board | Broadcom 57416 Dual Port 10GbE BASE-T & 5720 Dual Port 1GbE BASE-T, Network Daughter Card | Model: 1.2TB 10K RPM SAS 12Gbps 512n 2.5in Hot-plug Hard Drive No. of Drives – 3 Capacity – 3.6 TB |
| 2 | Dell PowerEdge T640 | Model: Intel Xeon Gold 6136 3.0G, 12C/24T, 10.4GT/s 2UPI, 24.75M Cache, Turbo, HT (150W) DDR4-2666, OEM XL No. Of CPUs: 1 | Model: 16GB RDIMM, 2666MT/s, Dual Rank Size: 32 GB | Default on Board | On board Broadcom 57416 Dual Port 10GbE LOM | Model: 1.2TB 10K RPM SAS 12Gbps 512n 2.5in Hot-plug Hard Drive, 3.5in HYB CARR No. of Drives – 3 Capacity – 3.6 TB |
| 3 | HP DL360 | Model (Both Proc): HP DL360 Gen9 E5-2643v No. Of CPUs: 2 | Model: HP 8GB (1x8GB) Single Rank x8 DDR4-2400 Size: 32 GB | Default on Board | HP Ethernet 1Gb 4-port 331i Adapter (Vendor: Broadcom) | Model: 1.2TB SAS 10K 2.5in No. of Drives – 3 Capacity – 3.6 TB |
| 4 | HP ML350 | Model (1st Proc): Intel Xeon E5-2643v4 3.4 GHz 6C Model (2nd Proc): HP ML350 Gen9 E5-2643v4 No. Of CPUs: 2 | Model: HP 8GB (1x8GB) Single Rank x8 DDR4-2400 Size: 32 GB | Default on Board | Embedded 1Gb 331i Ethernet Adapter 4 Ports per controller | Model: 1.2TB SAS 10K 2.5in No. of Drives – 3 Capacity – 3.6 TB |
| 5 | HP Z440 | Model: Intel Xeon E5-1660v4 3.20GHz 20MB 2400 8C CPU No. Of CPUs: 1 | Model: 16GB DDR4-2400 (2x8GB) ECC Reg RAM Size: 16 GB | NVIDIA Quadro K620 2GB DL-DVI(I)+DP | Default on Board | Model: 1TB 7200 RPM SATA 1st Hard Drive No. of Drives – 1 Capacity – 1 TB |
| 6 | HP Z4 (W-2145 Processor) | Model: Intel Xeon W-2145 3.7GHz 8C No. Of CPUs: 1 | Model: 16GB (2x8GB) DDR4 2666 DIMM ECC Registered Memory Size: 16 GB | NVIDIA Quadro P620 2GB (4) mDP | Default on Board | Model (1st Drive): HP Z Turbo Drive M.2 256GB SSD Model (2nd Drive): 1TB 7200RPM SATA 3.5in Enterprise No. of Drives – 2 Capacity – 1.2 TB |
| 7 | HP Z4 (W-2245 Processor) | Model: Intel Xeon W-2245 3.7GHz 8C No. Of CPUs: 1 | Model: 16GB (2x8GB) DDR4 2666 DIMM ECC Registered Memory Size: 16 GB | NVIDIA Quadro P620 2GB (4) mDP | Default on Board | Model (1st Drive): HP Z Turbo Drive M.2 256GB SSD Model (2nd Drive): 1TB 7200RPM SATA 3.5in Enterprise No. of Drives – 2 Capacity – 1.2 TB |
| 8 | HP Zbook G3 | Model: Xeon® E3-1545M v5 (2.9 GHz, 2133 MHz, 8 MB Cache, 4 cores) No. Of CPUs: 1 | Model: 64GB (4x16GB) 2133 DDR4 Size: 64 GB | NVIDIA Quadro M1000M 2GB GDDR5 Graphics | Default on Board | Model (1st Drive): M.2 SSD Storage 512GB Z Turbo Drive PCIe Solid State Drive Model (2nd Drive): drive M.2 SSD Storage 512GB Z Turbo Drive PCIe 2nd Solid State Drive No. of Drives – 2 Capacity – 1 TB |

###### Portal Client Prerequisites (Minimal Requirements Utility):

| Component | IntelliSpace Portal Client Specifications (Poor User Experience): |
| --- | --- |
| Supported OS | - Windows 10 64 bit - Recommended. - Windows 11 64 bit |
| Memory | - Minimum Memory (RAM): 4GB RAM. - Recommended: 8 GB RAM or above |
| Processor | - Minimum Processor (CPU) : 4 cores @ 2.4 GHz ; - Recommended: 6 Cores @ 3.0 GHz |
| Hard Drive | - Recommended Free Disk Space: 10 GB or above. - Additional 9 GB of free disk space are required to burn DVD. - Additional 15 GB of free disk space are required for iXR client |
| Network | - Minimum Network speed: 100 Mbps - Recommended: 1Gbps or above |
| Monitor | - Supported Screen Resolution: 1280x1024 up to 6MP. - Recommended: 1280x1024 (or above) - Layout: Landscape - Multi monitor support requires adequate support of client display card and driver |
| Software Pre-Requisites | - Net Framework 4.8 (Or later) - Web-Browser for PMT/CMT/ZFP/Client Download page: IE11, Chrome, Firefox, Microsoft Edge (latest version) (IE11 not supported post EOL) |
| Additional Software Recommended (for optional features) | - Adobe Acrobat Reader [for Report & Help] - Windows Media Player 9.0 or above [for saving Movies] - IMAPIv2 [for Burning CD/DVD] - DirectX 9.c for better CT/MR experience - Windows Management Instrumentation for MR & MMV applications - Microsoft Visual C++ 2013 (x64) 12.0.30501 for 4D Flow application - NVIDIA Grafics card with OpenGL 3.3, >=256 MB RAM for 4D Flow application |

###### Execution Plan on Supported Client Operating Systems:

|  | Windows 10 64 Bit |  |  | Windows 11 64 Bit |
| --- | --- | --- | --- | --- |
| Applications / Features | 20H2 (Build 19042) | 21H1 (Build 19043) | 21H2 (Build 19044) | 21H2 (Build 22000) |
| MR Cardiac Suite | X |  |  | X |
| MR Applications | X |  |  |  |
| CT Applications | X |  |  |  |
| System Features HFA | X |  |  |  |
| CT Cardiac Applications | X |  |  |  |
| CT Spectral Applications |  |  | X |  |
| CT Viewing |  |  |  | X |
| NM Applications | X | X | X | X |
| CT, MR, NM Legacy Applications | X |  |  | X |
| Installation, Upgrades & Security | X | X | X | X |
| System Features BLR | X |  |  | X |
| Serviceability | X |  |  | X |
| Non-Functional Tests |  |  | X |  |

###### Operating System & Database Requirements:

| Component | Operating System | Database |
| --- | --- | --- |
| Portal Client | Windows 10 (20H2, 21H1, 21H2 64 bit) Windows 11 (21H2 64 bit) | N/A |
|  | Citrix XenDesktop Virtual Desktop | N/A |
| Portal System | Windows Server 2019 [Standard and Data Center Editions] | MSSQL Express 2019 |
|  | Workstation - Windows 10 IoT Embedded LTSC OS (1809 Build) | MSSQL Express 2019 |
| Concerto System | Windows Server 2019 [Standard and Data Center Editions] | MSSQL 2019 Standard / Express |

###### Execution Plan on Deployments:

| Applications / Features | Demo Laptop | IX / LX (PIAB) | HX (Client Server) | EX (Master Slave) | EX – HA (Master Backup – High Availability) | Concerto |
| --- | --- | --- | --- | --- | --- | --- |
| MR Cardiac Suite | X | X | X | X |  | X |
| MR Applications |  |  | X |  |  |  |
| CT Applications |  | X | X |  |  | X |
| System Features HFA |  | X | X |  |  | X |
| CT Cardiac Applications |  | X | X | X |  | X |
| CT Spectral Applications |  | X | X | X |  |  |
| CT Viewing |  |  | X |  |  |  |
| NM Applications |  | X | X | X |  | X |
| CT, MR, NM Legacy Applications |  | X | X |  |  |  |
| Installation, Upgrades & Security | X | X | X | X | X | X |
| System Features BLR |  | X | X |  |  | X |
| Serviceability |  | X | X |  |  | X |
| Non-Functional Tests |  |  |  | X |  |  |

###### Execution Plan on Supported Reference Systems:

| Applications / Features | CT HAWK v5.0 | BB v4.8 | MR R5.9 | Vereos Digital PET/CT HeartBeat V2.0.3 | Forte JETStream | ISCV v6.2 | IBE B.18 | ISPACS v4.4.553 | ISPACS ISR 4.7.11 UDM 3.1.8 | VuePACS 12.2.8 | VuePACS PIM 15.0 | Allura 8.0.X | CardioMD IV - legacy | US: Affiniti 50C, Affiniti 70C, EPIQ 5C, EPIQ 7C, EPIQ CVx_9 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| MR Cardiac Suite |  |  | X |  |  |  |  |  |  | X | X |  |  |  |
| MR Applications |  |  | X |  |  |  |  |  |  |  |  |  |  |  |
| CT Applications | X | X |  |  |  |  |  |  |  |  |  |  |  |  |
| System Features HFA |  |  |  |  |  |  |  | X |  | X |  |  |  |  |
| CT Cardiac Applications | X | X |  |  |  |  |  |  |  |  |  |  |  |  |
| CT Spectral Applications | X |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NM Applications |  |  |  | X | X |  |  |  |  | X | X |  | X |  |
| CT, MR, NM Legacy Applications |  |  | X |  |  |  |  |  |  |  |  |  |  | X |
| System Features BLR |  |  |  |  |  | X | X | X | X | X | X | X |  |  |

##### AVW15 RC Release

###### Execution Plan on Supported Hardware:

|  | Client Server |  |  | PIAB (IX / LX) |  |  | Demo Laptop | Conc erto | Virtualization |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Applications / Features | DELL T/R 640 | HP DL-360 Gen9 (Rack) | HP ML-350 Gen9 (Tower) | HP Z440 | HP Z4 (W-2145 Processor) | HP Z4 (W-2245 Processor) | HP ZBook G3 | VMWare virtualization solution | VMware Virtual Machine (ESXi) | Microsoft Hyper V | Nutanix |
| MR Cardiac Suite | X | X |  | X |  |  | X | X |  |  |  |
| MR Applications |  | X |  |  | X |  |  | X |  |  |  |
| CT Applications | X |  |  |  |  | X |  | X | X |  |  |
| System Features HFA | X |  |  |  |  | X |  | X | X | X |  |
| CT Cardiac Applications | X | X | X | X | X | X | X | X |  |  |  |
| CT Spectral Applications |  |  | X | X |  |  |  |  |  |  |  |
| CT Viewing | X | X | X | X | X | X | X | X |  |  |  |
| NM Applications |  | X |  | X |  |  |  |  |  |  |  |
| CT, MR, NM Legacy Applications | X | X |  |  | X |  |  |  |  |  |  |
| Installation, Upgrades & Security | X | X | X | X | X | X | X | X | X | X | X |
| System Features BLR | X | X |  | X |  |  |  | X |  |  |  |
| Serviceability |  | X |  | X |  |  |  | X |  |  |  |
| Non-Functional Tests |  | X |  |  |  |  |  |  |  |  |  |

###### Hardware Specification:

| S.No | Hardware Model | CPU (Model and number of CPUs) | RAM (Model & Size) | Graphic Card (Where available) | NIC – Model of Network Interface Card | Storage: Model, No. of drives & capacity |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Dell PowerEdge R640 | Model: Intel Xeon Gold 6136 3.0G, 12C/24T, 10.4GT/s 2UPI, 24.75M Cache, Turbo, HT (150W) DDR4-2666, OEM XL No. Of CPUs: 1 | Model: 16GB RDIMM, 2666MT/s, Dual Rank Size: 32 GB | Default on Board | Broadcom 57416 Dual Port 10GbE BASE-T & 5720 Dual Port 1GbE BASE-T, Network Daughter Card | Model: 1.2TB 10K RPM SAS 12Gbps 512n 2.5in Hot-plug Hard Drive No. of Drives – 3 Capacity – 3.6 TB |
| 2 | Dell PowerEdge T640 | Model: Intel Xeon Gold 6136 3.0G, 12C/24T, 10.4GT/s 2UPI, 24.75M Cache, Turbo, HT (150W) DDR4-2666, OEM XL No. Of CPUs: 1 | Model: 16GB RDIMM, 2666MT/s, Dual Rank Size: 32 GB | Default on Board | On board Broadcom 57416 Dual Port 10GbE LOM | Model: 1.2TB 10K RPM SAS 12Gbps 512n 2.5in Hot-plug Hard Drive, 3.5in HYB CARR No. of Drives – 3 Capacity – 3.6 TB |
| 3 | HP DL360 | Model (Both Proc): HP DL360 Gen9 E5-2643v No. Of CPUs: 2 | Model: HP 8GB (1x8GB) Single Rank x8 DDR4-2400 Size: 32 GB | Default on Board | HP Ethernet 1Gb 4-port 331i Adapter (Vendor: Broadcom) | Model: 1.2TB SAS 10K 2.5in No. of Drives – 3 Capacity – 3.6 TB |
| 4 | HP ML350 | Model (1st Proc): Intel Xeon E5-2643v4 3.4 GHz 6C Model (2nd Proc): HP ML350 Gen9 E5-2643v4 No. Of CPUs: 2 | Model: HP 8GB (1x8GB) Single Rank x8 DDR4-2400 Size: 32 GB | Default on Board | Embedded 1Gb 331i Ethernet Adapter 4 Ports per controller | Model: 1.2TB SAS 10K 2.5in No. of Drives – 3 Capacity – 3.6 TB |
| 5 | HP Z440 | Model: Intel Xeon E5-1660v4 3.20GHz 20MB 2400 8C CPU No. Of CPUs: 1 | Model: 16GB DDR4-2400 (2x8GB) ECC Reg RAM Size: 16 GB | NVIDIA Quadro K620 2GB DL-DVI(I)+DP | Default on Board | Model: 1TB 7200 RPM SATA 1st Hard Drive No. of Drives – 1 Capacity – 1 TB |
| 6 | HP Z4 (W-2145 Processor) | Model: Intel Xeon W-2145 3.7GHz 8C No. Of CPUs: 1 | Model: 16GB (2x8GB) DDR4 2666 DIMM ECC Registered Memory Size: 16 GB | NVIDIA Quadro P620 2GB (4) mDP | Default on Board | Model (1st Drive): HP Z Turbo Drive M.2 256GB SSD Model (2nd Drive): 1TB 7200RPM SATA 3.5in Enterprise No. of Drives – 2 Capacity – 1.2 TB |
| 7 | HP Z4 (W-2245 Processor) | Model: Intel Xeon W-2245 3.7GHz 8C No. Of CPUs: 1 | Model: 16GB (2x8GB) DDR4 2666 DIMM ECC Registered Memory Size: 16 GB | NVIDIA Quadro P620 2GB (4) mDP | Default on Board | Model (1st Drive): HP Z Turbo Drive M.2 256GB SSD Model (2nd Drive): 1TB 7200RPM SATA 3.5in Enterprise No. of Drives – 2 Capacity – 1.2 TB |
| 8 | HP Zbook G3 | Model: Xeon® E3-1545M v5 (2.9 GHz, 2133 MHz, 8 MB Cache, 4 cores) No. Of CPUs: 1 | Model: 64GB (4x16GB) 2133 DDR4 Size: 64 GB | NVIDIA Quadro M1000M 2GB GDDR5 Graphics | Default on Board | Model (1st Drive): M.2 SSD Storage 512GB Z Turbo Drive PCIe Solid State Drive Model (2nd Drive): drive M.2 SSD Storage 512GB Z Turbo Drive PCIe 2nd Solid State Drive No. of Drives – 2 Capacity – 1 TB |

###### Portal Client Prerequisites (Minimal Requirements Utility):

| Component | IntelliSpace Portal Client Specifications (Poor User Experience): |
| --- | --- |
| Supported OS | - Windows 10 64 bit - Recommended. - Windows 11 64 bit |
| Memory | - Minimum Memory (RAM): 4GB RAM. - Recommended: 8 GB RAM or above |
| Processor | - Minimum Processor (CPU) : 4 cores @ 2.4 GHz ; - Recommended: 6 Cores @ 3.0 GHz |
| Hard Drive | - Recommended Free Disk Space: 10 GB or above. - Additional 9 GB of free disk space are required to burn DVD. - Additional 15 GB of free disk space are required for iXR client |
| Network | - Minimum Network speed: 100 Mbps - Recommended: 1Gbps or above |
| Monitor | - Supported Screen Resolution: 1280x1024 up to 6MP. - Recommended: 1280x1024 (or above) - Layout: Landscape - Multi monitor support requires adequate support of client display card and driver |
| Software Pre-Requisites | - Net Framework 4.8 (Or later) - Web-Browser for PMT/CMT/ZFP/Client Download page: IE11, Chrome, Firefox, Microsoft Edge (latest version) (IE11 not supported post EOL) |
| Additional Software Recommended (for optional features) | - Adobe Acrobat Reader [for Report & Help] - Windows Media Player 9.0 or above [for saving Movies] - IMAPIv2 [for Burning CD/DVD] - DirectX 9.c for better CT/MR experience - Windows Management Instrumentation for MR & MMV applications - Microsoft Visual C++ 2013 (x64) 12.0.30501 for 4D Flow application - NVIDIA Grafics card with OpenGL 3.3, >=256 MB RAM for 4D Flow application |

###### Execution Plan on Supported Client Operating Systems:

|  | Windows 10 64 Bit |  | Windows 11 64 Bit |  |
| --- | --- | --- | --- | --- |
| Applications / Features | 21H2 (Build 19044) | 22H2 (Build 19045) | 21H2(Build 22000) | 22H2 (Build 22621) |
| MR Cardiac Suite |  | X |  | X |
| MR Applications |  | X |  | X |
| CT Applications |  |  |  |  |
| System Features HFA |  |  |  |  |
| CT Cardiac Applications | X |  | X |  |
| CT Spectral Applications |  | X |  | X |
| CT Viewing | X |  | X |  |
| NM Applications | X | X | X | X |
| CT, MR, NM Legacy Applications | X |  |  | X |
| Installation, Upgrades & Security | X | X | X | X |
| System Features BLR | X |  | X |  |
| Serviceability | NA | NA | NA | NA |
| Non-Functional Tests |  | X |  |  |

###### Operating System & Database Requirements:

| Component | Operating System | Database |
| --- | --- | --- |
| Portal Client | Windows 10 (21H2,22H2 64 bit) Windows 11 (21H2,22H2 64 bit) | N/A |
|  | Citrix XenDesktop Virtual Desktop | N/A |
| Portal System | Windows Server 2019 [Standard and Data Center Editions] | MSSQL Express 2019 |
|  | Workstation - Windows 10 IoT Embedded LTSC OS (1809 Build) | MSSQL Express 2019 |
| Concerto System | Windows Server 2019 [Standard and Data Center Editions] | MSSQL 2019 Standard / Express |

###### Execution Plan on Deployments:

| Applications / Features | Demo Laptop | IX / LX (PIAB) | HX (Client Server) | EX (Master Slave) | EX – HA (Master Backup – High Availability) | Concerto |
| --- | --- | --- | --- | --- | --- | --- |
| MR Cardiac Suite | X | X | X | X | X | X |
| MR Applications |  | X | X | X |  |  |
| CT Applications |  | X | X |  |  | X |
| System Features HFA |  | X | X |  |  | X |
| CT Cardiac Applications |  | X | X | X |  | X |
| CT Spectral Applications |  | X | X | X |  |  |
| CT Viewing |  |  | X |  |  |  |
| NM Applications |  | X | X | X |  | X |
| CT, MR, NM Legacy Applications |  | X | X |  |  |  |
| Installation, Upgrades & Security | X | X | X | X | X | X |
| System Features BLR |  | X | X |  |  | X |
| Serviceability |  | X | X |  |  | X |
| Non-Functional Tests |  |  | X |  |  |  |

###### Execution Plan on Supported Reference Systems:

| Applications / Features | CT Incisive 1.1 HB V4.5.1.742 | ISCV v7.0 | CT Incisive for Dual Energy verification per marketing | Allura 8. x.100 | US: More legacy cameras if the data is available |
| --- | --- | --- | --- | --- | --- |
| MR Cardiac Suite |  |  |  |  |  |
| MR Applications |  |  |  |  |  |
| CT Applications | X |  | X |  |  |
| System Features HFA |  |  |  |  |  |
| CT Cardiac Applications |  |  |  |  |  |
| CT Spectral Applications |  |  |  |  |  |
| CT Viewing |  |  |  |  |  |
| NM Applications |  |  |  |  |  |
| CT, MR, NM Legacy Applications |  |  |  |  | X |
| Installation, Upgrades & Security |  |  |  |  |  |
| System Features BLR |  | X |  | X |  |
| Serviceability |  |  |  |  |  |

### CT ASPECT Scoring (CT ASPECTS)

#### Target hardware

##### RFD Release:

###### Execution Plan on Supported Hardware:

|  | Client Server |  |  | PIAB (IX / LX) |  |  | Demo Laptop | Conc erto | Virtualization |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Applications / Features | DELL T/R 640 | HP DL-360 Gen9 (Rack) | HP ML-350 Gen9 (Tower) | HP Z440 | HP Z4 (W-2145 Processor) | HP Z4 (W-2245 Processor) | HP ZBook G3 | VMWare virtualization solution | VMware Virtual Machine (ESXi) | Microsoft Hyper V | Nutanix |
| CT ASPECT Scoring (CT ASPECTS) |  |  |  |  |  |  |  |  | X |  |  |

###### Execution Plan on Supported Client Operating Systems:

|  | Windows 10 64 Bit |  |  | Windows 11 64 Bit |
| --- | --- | --- | --- | --- |
| Applications / Features | 20H2 (Build 19042) | 21H1 (Build 19043) | 21H2 (Build 19044) | 21H2 (Build 22000) |
| CT ASPECT Scoring (CT ASPECTS) | X |  |  |  |

###### Execution Plan on Deployments:

| Applications / Features | Demo Laptop | IX / LX (PIAB) | HX (Client Server) | EX (Master Slave) | EX – HA (Master Backup – High Availability) | Concerto |
| --- | --- | --- | --- | --- | --- | --- |
| CT ASPECT Scoring (CT ASPECTS) |  |  | X |  |  |  |

##### RC Release:

###### Execution Plan on Supported Hardware:

|  | Client Server |  |  | PIAB (IX / LX) |  |  | Demo Laptop | Conc erto | Virtualization |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Applications / Features | DELL T/R 640 | HP DL-360 Gen9 (Rack) | HP ML-350 Gen9 (Tower) | HP Z440 | HP Z4 (W-2145 Processor) | HP Z4 (W-2245 Processor) | HP ZBook G3 | VMWare virtualization solution | VMware Virtual Machine (ESXi) | Microsoft Hyper V | Nutanix |
| CT ASPECT Scoring (CT ASPECTS) |  |  |  |  |  |  |  |  | X |  |  |

###### Execution Plan on Supported Client Operating Systems:

|  | Windows 10 64 Bit |  | Windows 11 64 Bit |  |
| --- | --- | --- | --- | --- |
| Applications / Features | 21H2 (Build 19044) | 22H2 (Build 19045) | 21H2(Build 22000) | 22H2 (Build 22621) |
| CT ASPECT Scoring (CT ASPECTS) | X |  |  |  |

###### Execution Plan on Deployments:

| Applications / Features | Demo Laptop | IX / LX (PIAB) | HX (Client Server) | EX (Master Slave) | EX – HA (Master Backup – High Availability) | Concerto |
| --- | --- | --- | --- | --- | --- | --- |
| CT ASPECT Scoring (CT ASPECTS) |  |  | X |  |  |  |

### NM Suite

#### Target hardware

##### RFD Release:

###### Execution Plan on Supported Hardware:

|  | Client Server |  |  | PIAB (IX / LX) |  |  | Demo Laptop | Conc erto | Virtualization |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Applications / Features | DELL T/R 640 | HP DL-360 Gen9 (Rack) | HP ML-350 Gen9 (Tower) | HP Z440 | HP Z4 (W-2145 Processor) | HP Z4 (W-2245 Processor) | HP ZBook G3 | VMWare virtualization solution | VMware Virtual Machine (ESXi) | Microsoft Hyper V | Nutanix |
| NM Suite |  | X |  | X |  |  |  |  |  |  |  |

###### Execution Plan on Supported Client Operating Systems:

|  | Windows 10 64 Bit |  |  | Windows 11 64 Bit |
| --- | --- | --- | --- | --- |
| Applications / Features | 20H2 (Build 19042) | 21H1 (Build 19043) | 21H2 (Build 19044) | 21H2 (Build 22000) |
| NM Suite | X | X | X | X |

###### Execution Plan on Deployments:

| Applications / Features | Demo Laptop | IX / LX (PIAB) | HX (Client Server) | EX (Master Slave) | EX – HA (Master Backup – High Availability) | Concerto |
| --- | --- | --- | --- | --- | --- | --- |
| NM Suite |  | X | X |  |  |  |

##### RC Release

###### Execution Plan on Supported Hardware:

|  | Client Server |  |  | PIAB (IX / LX) |  |  | Demo Laptop | Conc erto | Virtualization |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Applications / Features | DELL T/R 640 | HP DL-360 Gen9 (Rack) | HP ML-350 Gen9 (Tower) | HP Z440 | HP Z4 (W-2145 Processor) | HP Z4 (W-2245 Processor) | HP ZBook G3 | VMWare virtualization solution | VMware Virtual Machine (ESXi) | Microsoft Hyper V | Nutanix |
| NM Suite |  | X |  | X |  |  |  |  |  |  |  |

###### Execution Plan on Supported Client Operating Systems:

|  | Windows 10 64 Bit |  | Windows 11 64 Bit |  |
| --- | --- | --- | --- | --- |
| Applications / Features | 21H2 (Build 19044) | 22H2 (Build 19045) | 21H2(Build 22000) | 22H2 (Build 22621) |
| NM Suite | X | X | X | X |

###### Execution Plan on Deployments:

| Applications / Features | Demo Laptop | IX / LX (PIAB) | HX (Client Server) | EX (Master Slave) | EX – HA (Master Backup – High Availability) | Concerto |
| --- | --- | --- | --- | --- | --- | --- |
| NM Suite |  | X | X |  |  |  |

## Product Hazard/Risk Management Activities

A project’s Risk Management Files have been prepared to identify, track and mitigate all issues determined as potential hazards to patient, operator or service personnel, and assigned all relevant hazards to all sub-teams of IntelliSpace Portal.

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release:

| DHF No | Document Name | Version |
| --- | --- | --- |
| D000991970 | ISP13 RMA MR Cardiac Suite | A |
| D000879224 | ISP13 Risk Management Matrix Cardiac Suite | A |
| D000938106 | ISP13 RMA for CTApps | C |
| D000892550 | ISP13 RMM General CTApps | C |
| D000946872 | ISP13 RMA CT Cardiac | A |
| D000946860 | ISP13_Risk Management Matrix_CT Cardiac | A |

RMM Tests Execution in ISP13 VAL-3 Development Phase:

| DHF No | Document Name | Version |
| --- | --- | --- |
| D001065109 | ISP13_Risk Management Analysis_MRApps & MMV | A |
| D001065108 | ISP13_RMM_MRAPPS & MMV | A |
| D001065179 | ISP13 Legacy MR Cardiac RMA | A |
| D001064798 | ISP13 MR Cardiac Legacy applications Hazard Test Spec | A |
| D001059457 | ISP13_NM Apps_RMA | A |
| D001059463 | ISP13_NM Apps_RMM | A |
| D000938106 | ISP13 RMA for CTApps | C |
| D000892550 | ISP13 RMM General CTApps | C |
| D000946872 | ISP13 RMA CT Cardiac | A |
| D000946860 | ISP13_Risk Management Matrix_CT Cardiac | A |
| D000952406 | ISP13 RMA CTSpectral Apps | A |
| D000996304 | ISP13 RMM CTSpectral Apps | A |

#### ISP13 VAL-3 Release:

| DHF No | Document Name | Version |
| --- | --- | --- |
| D001065109 | ISP13_Risk Management Analysis_MRApps & MMV | A |
| D001065108 | ISP13_RMM_MRAPPS & MMV | A |
| D001065179 | ISP13 Legacy MR Cardiac RMA | A |
| D001064798 | ISP13 MR Cardiac Legacy applications Hazard Test Spec | A |
| D000991970 | ISP13 RMA MR Cardiac Suite | A |
| D000879224 | ISP13 Risk Management Matrix Cardiac Suite | B |
| D001105931 | ISP13 LoBI RMA | A |
| D001104871 | ISP13 LoBI RMM | A |
| D001059457 | ISP13_NM Apps_RMA | A |
| D001059463 | ISP13_NM Apps_RMM | A |
| D000946872 | ISP13 RMA CT Cardiac | A |
| D000946860 | ISP13_Risk Management Matrix_CT Cardiac | A |
| D000938106 | ISP13 RMA for CTApps | D |
| D000892550 | ISP13 RMM General CTApps | D |
| D000952406 | ISP13 RMA CTSpectral Apps | B |
| D000996304 | ISP13 RMM CTSpectral Apps | B |
| D000996286 | ISP13 System RMM | A |

RMM Tests Execution in AVW15 RFD Development Phase:

| DHF No | Document Name | Version |
| --- | --- | --- |
| D000991970 | ISP13 RMA MRCardiac Suite | B |
| D000879224 | ISP13 RISK MANAGEMENT MATRIX CARDIAC SUITE | C |
| D001065109 | ISP13_Risk Management Analysis_MRApps & MMV | A |
| D001105931 | ISP13 RMA LoBI | A |
| D001104871 | Risk Management Matrix for ISP13,MR Applications and MMV | B |
| D001065108 | ISP13_RMM_MRAPPS & MMV | A |
| D000938106 | ISP13 RMA for CTapps | D |
| D000892550 | ISP13 RMM General CTApps | D |
| D000946872 | ISP13 RMA CT Cardiac | A |
| D000946860 | ISP13 RMM CT Cardiac | A |
| D000952406 | ISP13 RMA CTSpectral Apps | B |
| D000996304 | ISP13 RMM CTSpectral Apps | B |
| D001059457 | ISP13 NM Apps_RMA | A |
| D001059463 | ISP13 NM Apps_RMM | A |
| DHF380158 | ISP13 RMA for ISP System | LC |
| D000996286 | ISP13 System RMM | A |

#### AVW15 RFD Release:

| DHF No | Document Name | Version |
| --- | --- | --- |
| D000991970 | ISP13 RMA MR Cardiac Suite | C |
| D000879224 | ISP13 Risk Management Matrix Cardiac Suite | D |
| D001065108 | ISP13_RMM_MRAPPS & MMV | A |
| D001065109 | ISP13_Risk Management Analysis_MRApps & MMV | A |
| D000938106 | ISP13 RMA for CTapps | E |
| D000892550 | ISP13 RMM General CTApps | E |
| D000946860 | ISP13_Risk Management Matrix_CT Cardiac | B |
| D000946872 | ISP13 RMA CT Cardiac | B |
| D000952406 | ISP13 RMA CTSpectral Apps | C |
| D000996304 | ISP13 RMM CTSpectral Apps | C |
| D000996286 | ISP13 System RMM | B |

#### AVW15 RC Release

| DHF No | Document Name | Version |
| --- | --- | --- |
| D000991970 | ISP13 RMA MR Cardiac Suite | C |
| D000879224 | ISP13 Risk Management Matrix Cardiac Suite | D |
| D001065108 | ISP13_RMM_MRAPPS & MMV | A |
| D001065109 | ISP13_Risk Management Analysis_MRApps & MMV | A |
| D000938106 | ISP13 RMA for CTapps | G |
| D000892550 | ISP13 RMM General CTApps | G |
| D000946860 | ISP13_Risk Management Matrix_CT Cardiac | B |
| D000946872 | ISP13 RMA CT Cardiac | B |
| D000952406 | ISP13 RMA CTSpectral Apps | C |
| D000996304 | ISP13 RMM CTSpectral Apps | C |
| D000996286 | ISP13 System RMM | B |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release:

| DHF No | Document Name | Version |
| --- | --- | --- |
| D001157471 | RMM General CTApps | A |

#### RC Release:

| DHF No | Document Name | Version |
| --- | --- | --- |
| D001157471 | RMM General CTApps | A |

### NM Suite

#### RFD Release:

| DHF No | Document Name | Version |
| --- | --- | --- |
| D001024262 | ISP12.4 _RMA_NM Apps | A |
| D001024288 | ISP12.4_NM Apps_RMM | A |

#### RC Release:

| DHF No | Document Name | Version |
| --- | --- | --- |
| D001024262 | ISP12.4 _RMA_NM Apps | A |
| D001024288 | ISP12.4_NM Apps_RMM | A |

## Exploratory Testing Strategy

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

NA, No Exploratory tests are planned for ISP13 VAL-2 release.

#### ISP13 VAL-3 Release

NA, No Exploratory tests are planned for ISP13 VAL-3 release.

#### AVW15 RFD Release

NA, No Exploratory tests are planned for AVW15 RFD release.

#### AVW15 RC Release

NA, No Exploratory tests are planned for AVW15 RC release.

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

NA, No Exploratory tests are planned for RFD release.

#### RC Release

NA, No Exploratory tests are planned for RC release.

### NM Suite

#### RFD Release

NA, No Exploratory tests are planned for RFD release.

#### RC Release

NA, No Exploratory tests are planned for RC release.

## Interface Testing Strategy

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

NA, No Interface tests are planned for ISP13 VAL-2 release.

#### ISP13 VAL-3 Release

NA, No Interface tests are planned for ISP13 VAL-3 release.

#### AVW15 RFD Release

NA, No Interface tests are planned for AVW15 RFD release.

#### AVW15 RC Release

NA, No Interface tests are planned for AVW15 RC release.

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

NA, No Interface tests are planned for RFD release.

#### RC Release

NA, No Interface tests are planned for RC release.

### NM Suite

#### RFD Release

NA, No Interface tests are planned for RFD release.

#### RC Release

NA, No Interface tests are planned for RC release.

## Usability Engineering - Formative Evaluation Strategy

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

DHF380228 ISP13 Usability Engineering Plan Rev LA

#### ISP13 VAL-3 Release

DHF380228 ISP13 Usability Engineering Plan Rev LB

#### AVW15 RC Release

DHF380228 ISP13 Usability Engineering Plan Rev LE

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

DHF380228 ISP13 Usability Engineering Plan Rev LB

#### RC Release

DHF380228 ISP13 Usability Engineering Plan Rev LE

### NM Suite

#### RFD Release

DHF380228 ISP13 Usability Engineering Plan Rev LB

#### RC Release

DHF380228 ISP13 Usability Engineering Plan Rev LE

# Resources

Execution of the scope mentioned in section 3.2.1 (Coverage Strategy) will be carried at sub-team level. Resources with the role and sub-team detail is mentioned in below table.

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

| Name | Role | Modalities | Tasks |
| --- | --- | --- | --- |
| Srinivas Arakere | Global Verification Manager | N/A | Execution Tracking, Status Reporting for ISP. |
| Aashish Kumar | Global Verification Lead | N/A | Execution Tracking, Status Reporting for ISP. |
| Mustaq Mohammed | Verification Lead | MR Cardiac Suite | Impact Analysis, Scope Sharing & Execution |
| Barbarash Vladimir | Verification Lead | CT Apps – CT Liver | Impact Analysis, Scope Sharing & Execution |
| Raju Gopi | Verification Lead | CT Apps – CCA | Impact Analysis, Scope Sharing & Execution |
| Lokesh R | Verification Engineer | MR Cardiac Suite | Execution |
| Praveenkumar SB | Verification Engineer | MR Cardiac Suite | Execution |
| Saar Hen | Verification Engineer | MR Cardiac Suite | Execution |
| Vijeta SM | Verification Engineer | MR Cardiac Suite | Execution |
| Archana Bhat | Verification Engineer | MR Cardiac Suite | Execution |
| Divya MR | Verification Engineer | MR Cardiac Suite | Execution |
| Sevinj Yusifova | Verification Engineer | CT Liver | Execution |
| Zehavit Ganon | Verification Engineer | CT Liver | Execution |
| Odeliya Klochko | Verification Engineer | CT Liver | Execution |
| Cristina Munteanu | Verification Engineer | CT Liver | Execution |
| Chandrashekhar Venkatraman | Verification Engineer | CCA | Execution |
| Neelam Hembrom | Verification Engineer | CCA | Execution |

RMM Tests Execution in ISP13 VAL-3 Development Phase:

| Name | Role | Modalities | Tasks |
| --- | --- | --- | --- |
| Srinivas Arakere | Global Verification Manager | N/A | Execution Tracking, Status Reporting for ISP. |
| Aashish Kumar | Global Verification Lead | N/A | Execution Tracking, Status Reporting for ISP. |
| Santhosh H | Verification Lead | MR / NM / CT Apps – Legacy Applications | Impact Analysis, Scope Sharing & Execution |
| Rudravarapu Kasiviswanath | Verification Lead | NM Apps | Impact Analysis, Scope Sharing & Execution |
| Barbarash Vladimir | Verification Lead | CT Apps | Impact Analysis, Scope Sharing & Execution |
| Raju Gopi | Verification Lead | CT Apps – CCA | Impact Analysis, Scope Sharing & Execution |
| Eti Shohat | Verification Lead | CT Spectral Apps | Impact Analysis, Scope Sharing & Execution |
| Chandrashekhar Venkatraman | Verification Lead | CT Viewing | Impact Analysis, Scope Sharing & Execution |
| Arthi Deenadayalan | Verification Engineer | MR / NM / CT Apps – Legacy Applications | Execution |
| Divya MR | Verification Engineer | MR / NM / CT Apps – Legacy Applications | Execution |
| Neelam Hembrom | Verification Engineer | MR / NM / CT Apps – Legacy Applications | Execution |
| Haindavi Karanam | Verification Engineer | MR / NM / CT Apps – Legacy Applications | Execution |
| Sevinj Yusifova | Verification Engineer | CT Apps | Execution |
| Zehavit Ganon | Verification Engineer | CT Apps | Execution |
| Odeliya Klochko | Verification Engineer | CT Apps | Execution |
| Cristina Munteanu | Verification Engineer | CT Apps | Execution |
| Elad Alexander | Verification Engineer | CT Spectral Apps | Execution |
| Danny Ben-yakar | Verification Engineer | CT Spectral Apps | Execution |
| Hagit Eizen | Verification Engineer | CT Spectral Apps | Execution |
| Liad Siles | Verification Engineer | CT Spectral Apps | Execution |
| Olga Treibich | Verification Engineer | CT Spectral Apps | Execution |
| Lilia Yermakov | Verification Engineer | CT Spectral Apps | Execution |
| Reut Zur | Verification Engineer | CT Spectral Apps | Execution |

#### ISP13 VAL-3 Release

| Name | Role | Modalities | Tasks |
| --- | --- | --- | --- |
| Srinivas Arakere | Global Verification Manager | N/A | Execution Tracking, Status Reporting for ISP. |
| Aashish Kumar | Global Verification Lead | N/A | Execution Tracking, Status Reporting for ISP. |
| Lokesh R | Verification Lead | MR Cardiac Suite + MR Apps | Impact Analysis, Scope Sharing & Execution |
| Raju Gopi | Verification Lead | CT Cardiac Apps | Impact Analysis, Scope Sharing & Execution |
| Barbarash Vladimir | Verification Lead | CT Apps – sAVA & Auto CS + System Features (HFA) | Impact Analysis, Scope Sharing & Execution |
| Eti Shohat | Verification Lead | CT Spectral Apps | Impact Analysis, Scope Sharing & Execution |
| Santhosh H | Verification Lead | MR / NM / CT Apps – Legacy Applications + Installation | Impact Analysis, Scope Sharing & Execution |
| Archana Shetty | Verification Lead | System Features (BLR) | Impact Analysis, Scope Sharing & Execution |
| Deepika Anand Hegde | Verification Engineer | MR Cardiac Suite + MR Apps | Execution |
| Rudravarapu Kasiviswanath | Verification Engineer | MR Cardiac Suite + NM Apps | Execution |
| Praveenkumar SB | Verification Engineer | MR Cardiac Suite + MR Apps | Execution |
| Chandrashekhar Venkatraman | Verification Engineer | MR Cardiac Suite | Execution |
| Laxmish Kulkarni | Verification Engineer | MR Cardiac Suite + MR Apps | Execution |
| Klochko Odeliya | Verification Engineer | CT Apps | Execution |
| Yusifova Sevinj | Verification Engineer | CT Apps | Execution |
| Munteanu Cristina | Verification Engineer | CT Apps | Execution |
| Ganon Zehavit | Verification Engineer | CT Apps | Execution |
| Markreich Gady | Verification Engineer | System Features (HFA) | Execution |
| Seligman Tania | Verification Engineer | System Features (HFA) | Execution |
| Elad Alexander | Verification Engineer | CT Spectral Apps | Execution |
| Danny Ben-yakar | Verification Engineer | CT Spectral Apps | Execution |
| Hagit Eizen | Verification Engineer | CT Spectral Apps | Execution |
| Lilia Yermakov | Verification Engineer | CT Spectral Apps | Execution |
| Reut Zur | Verification Engineer | CT Spectral Apps | Execution |
| Arthi Deenadayalan | Verification Engineer | MR / NM / CT Apps – Legacy Applications | Execution |
| Divya MR | Verification Engineer | MR / NM / CT Apps – Legacy Applications + Installation | Execution |
| Haindavi Karanam | Verification Engineer | MR / NM / CT Apps – Legacy Applications | Execution |
| Kavya Vishwakumar | Verification Engineer | System Features (BLR) | Execution |
| Anil Kumar | Verification Engineer | System Features (BLR) | Execution |
| Ganesh Srinivasarao Balijepalli | Verification Engineer | System Features (BLR) | Execution |

RMM Tests Execution in AVW15 RFD Development Phase:

| Name | Role | Modalities | Tasks |
| --- | --- | --- | --- |
| Srinivas Arakere | Global Verification Manager | N/A | Execution Tracking, Status Reporting for ISP. |
| Aashish Kumar | Global Verification Lead | N/A | Execution Tracking, Status Reporting for ISP. |
| Lokesh BR | Verification Lead | MR / MR Cardiac Apps | Impact Analysis, Scope Sharing & Execution |
| Barbarash Vladimir | Verification Lead | CT Apps | Impact Analysis, Scope Sharing & Execution |
| Raju Gopi | Verification Lead | CT Cardiac Apps | Impact Analysis, Scope Sharing & Execution |
| Eti Shohat | Verification Lead | CT Spectral Apps | Impact Analysis, Scope Sharing & Execution |
| Chandrashekhar Venkatraman | Verification Lead | CT Viewing | Impact Analysis, Scope Sharing & Execution |
| Rudravarapu Kasiviswanath | Verification Lead | NM Apps | Impact Analysis, Scope Sharing & Execution |
| Santhosh H | Verification Lead | MR / NM / CT Apps – Legacy Applications | Impact Analysis, Scope Sharing & Execution |
| Hegde Deepika Anand | Verification Engineer | MR / MR Cardiac Apps | Execution |
| Kulkarni Laxmish | Verification Engineer | MR / MR Cardiac Apps | Execution |
| S B Praveenkumar | Verification Engineer | MR / MR Cardiac Apps | Execution |
| Sevinj Yusifova | Verification Engineer | CT Apps | Execution |
| Zehavit Ganon | Verification Engineer | CT Apps | Execution |
| Odeliya Klochko | Verification Engineer | CT Apps | Execution |
| Cristina Munteanu | Verification Engineer | CT Apps | Execution |
| Elad Alexander | Verification Engineer | CT Spectral Apps | Execution |
| Danny Ben-yakar | Verification Engineer | CT Spectral Apps | Execution |
| Hagit Eizen | Verification Engineer | CT Spectral Apps | Execution |
| Reut Zur | Verification Engineer | CT Spectral Apps | Execution |
| Lilia Yermakov | Verification Engineer | CT Spectral Apps | Execution |
| Arthi Deenadayalan | Verification Engineer | MR / NM / CT Apps – Legacy Applications | Execution |
| Divya MR | Verification Engineer | MR / NM / CT Apps – Legacy Applications | Execution |
| Haindavi Karanam | Verification Engineer | MR / NM / CT Apps – Legacy Applications | Execution |
| Kavya Vishwakumar | Verification Engineer | System Features BLR | Execution |
| Ganesh Balijepalli | Verification Engineer | System Features BLR | Execution |
| Anil Kumar | Verification Engineer | System Features BLR | Execution |

#### AVW15 RFD Release

| Name | Role | Modalities | Tasks |
| --- | --- | --- | --- |
| Srinivas Arakere | Global Verification Manager | N/A | Execution Tracking, Status Reporting for ISP. |
| Aashish Kumar | Global Verification Lead | N/A | Execution Tracking, Status Reporting for ISP. |
| Lokesh R | Verification Lead | MR Cardiac Suite + MR Applications | Impact Analysis, Scope Sharing & Execution |
| Barbarash Vladimir | Verification Lead | CT Applications + System Features (HFA) | Impact Analysis, Scope Sharing & Execution |
| Raju Gopi | Verification Lead | CT Cardiac Applications | Impact Analysis, Scope Sharing & Execution |
| Eti Shohat | Verification Lead | CT Spectral Applications | Impact Analysis, Scope Sharing & Execution |
| Chandrashekhar Venkatraman | Verification Lead | CT Viewing + CT Cardiac Applications | Impact Analysis, Scope Sharing & Execution |
| Rudravarapu Kasiviswanath | Verification Lead | NM Applications + MR Cardiac Suite | Impact Analysis, Scope Sharing & Execution |
| Arthi D | Verification Lead | CT, MR, NM Legacy Applications | Impact Analysis, Scope Sharing & Execution |
| Santhosh H | Verification Lead | Installation, Upgrades & Security | Impact Analysis, Scope Sharing & Execution |
| Archana Shetty | Verification Lead | System Features BLR + Serviceability | Impact Analysis, Scope Sharing & Execution |
| Kavya Vishwakumar | Verification Lead | System Features BLR + Serviceability | Impact Analysis, Scope Sharing & Execution |
| Gaurav Sharma | Verification Lead | Non-Functional Tests | Impact Analysis, Scope Sharing & Execution |
| Deepika Anand Hegde | Verification Engineer | MR Cardiac Suite + MR Applications | Execution |
| Praveenkumar SB | Verification Engineer | MR Cardiac Suite + MR Applications | Execution |
| Klochko Odeliya | Verification Engineer | CT Applications | Execution |
| Yusifova Sevinj | Verification Engineer | CT Applications | Execution |
| Munteanu Cristina | Verification Engineer | CT Applications | Execution |
| Ganon Zehavit | Verification Engineer | CT Applications | Execution |
| Markreich Gady | Verification Engineer | System Features HFA | Execution |
| Seligman Tania | Verification Engineer | System Features HFA | Execution |
| Laxmish Kulkarni | Verification Engineer | CT Cardiac Applications | Execution |
| Bharath MR | Verification Engineer | CT Cardiac Applications | Execution |
| Elad Alexander | Verification Engineer | CT Spectral Applications | Execution |
| Danny Ben-yakar | Verification Engineer | CT Spectral Applications | Execution |
| Hagit Eizen | Verification Engineer | CT Spectral Applications | Execution |
| Lilia Yermakov | Verification Engineer | CT Spectral Applications | Execution |
| Reut Zur | Verification Engineer | CT Spectral Applications | Execution |
| Shakti Prasad Bavoor | Verification Engineer | NM Applications | Execution |
| Arthi Deenadayalan | Verification Engineer | CT, MR, NM Legacy Applications | Execution |
| Divya MR | Verification Engineer | CT, MR, NM Legacy Applications + Installation | Execution |
| Haindavi Karanam | Verification Engineer | CT, MR, NM Legacy Applications | Execution |
| Kothika Sarkar | Verification Engineer | CT, MR, NM Legacy Applications | Execution |
| Kunal Tomar | Verification Engineer | CT, MR, NM Legacy Applications + Installation | Execution |
| Anil Kumar | Verification Engineer | System Features BLR + Serviceability | Execution |
| Ganesh Srinivasarao Balijepalli | Verification Engineer | System Features BLR + Serviceability | Execution |
| Mohammed Zabi Pasha | Verification Engineer | Non-Functional Tests | Execution |
| Shruthi P | Verification Engineer | Non-Functional Tests | Execution |

#### AVW15 RC Release

| Name | Role | Modalities | Tasks |
| --- | --- | --- | --- |
| Srinivas Arakere | Global Verification Manager | N/A | Execution Tracking, Status Reporting for ISP. |
| Santhosh H | Global Verification Lead | N/A | Execution Tracking, Status Reporting for ISP. |
| Lokesh R | Verification Lead | MR Cardiac Suite + MR Applications | Impact Analysis, Scope Sharing & Execution |
| Barbarash Vladimir | Verification Lead | CT Applications + System Features (HFA) | Impact Analysis, Scope Sharing & Execution |
| Chandrashekhar Venkatraman | Verification Lead | CT Cardiac Applications | Impact Analysis, Scope Sharing & Execution |
| Haindavi Karanam | Verification Lead | CT Spectral Applications | Impact Analysis, Scope Sharing & Execution |
| Chandrashekhar Venkatraman | Verification Lead | CT Viewing + CT Cardiac Applications | Impact Analysis, Scope Sharing & Execution |
| Rudravarapu Kasiviswanath | Verification Lead | NM Applications + MR Cardiac Suite | Impact Analysis, Scope Sharing & Execution |
| Arthi D | Verification Lead | CT, MR, NM Legacy Applications | Impact Analysis, Scope Sharing & Execution |
| Santhosh H | Verification Lead | Installation, Upgrades & Security | Impact Analysis, Scope Sharing & Execution |
| Archana Shetty | Verification Lead | System Features BLR + Serviceability | Impact Analysis, Scope Sharing & Execution |
| Kavya Vishwakumar | Verification Lead | System Features BLR + Serviceability | Impact Analysis, Scope Sharing & Execution |
| Kiran B | Verification Lead | Non-Functional Tests | Impact Analysis, Scope Sharing & Execution |
| Deepika Anand Hegde | Verification Engineer | MR Cardiac Suite + MR Applications | Execution |
| Praveenkumar SB | Verification Engineer | MR Cardiac Suite + MR Applications | Execution |
| Klochko Odeliya | Verification Engineer | CT Applications | Execution |
| Nadav Judith | Verification Engineer | CT Applications | Execution |
| Markreich Gady | Verification Engineer | System Features HFA | Execution |
| Laxmish Kulkarni | Verification Engineer | CT Cardiac Applications | Execution |
| Shakti Prasad Bavoor | Verification Engineer | NM Applications | Execution |
| Divya MR | Verification Engineer | CT, MR, NM Legacy Applications + Installation | Execution |
| Haindavi Karanam | Verification Engineer | CT, MR, NM Legacy Applications and CT Spectral Applications | Execution |
| Kothika Sarkar | Verification Engineer | CT, MR, NM Legacy Applications and CT Spectral Applications | Execution |
| Kunal Tomar | Verification Engineer | CT, MR, NM Legacy Applications + Installation | Execution |
| Anil Kumar | Verification Engineer | System Features BLR + Serviceability | Execution |
| Ganesh Srinivasarao Balijepalli | Verification Engineer | System Features BLR + Serviceability | Execution |
| Mohammed Zabi Pasha | Verification Engineer | Non-Functional Tests | Execution |
| Shruthi P | Verification Engineer | Non-Functional Tests | Execution |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

| Name | Role | Modalities | Tasks |
| --- | --- | --- | --- |
| Aashish Kumar | Global Verification Lead | N/A | Execution Tracking, Status Reporting for ISP. |
| Barbarash Vladimir | Verification Lead | CT Applications – ASPECTS | Impact Analysis, Scope Sharing & Execution |
| Munteanu Cristina | Verification Engineer | CT Applications – ASPECTS | Execution |

#### RC Release

| Name | Role | Modalities | Tasks |
| --- | --- | --- | --- |
| Santhosh H | Global Verification Lead | N/A | Execution Tracking, Status Reporting for ISP. |
| Barbarash Vladimir | Verification Lead | CT Applications – ASPECTS | Impact Analysis, Scope Sharing & Execution |
| Munteanu Cristina | Verification Engineer | CT Applications – ASPECTS | Execution |

### NM Suite

#### RFD Release

| Name | Role | Modalities | Tasks |
| --- | --- | --- | --- |
| Rudravarapu Kasiviswanath | Verification Lead | NM Applications | Impact Analysis, Scope Sharing & Execution |
| Shakti Prasad Bavoor | Verification Engineer | NM Applications | Execution |

#### RC Release

| Name | Role | Modalities | Tasks |
| --- | --- | --- | --- |
| Rudravarapu Kasiviswanath | Verification Lead | NM Applications | Impact Analysis, Scope Sharing & Execution |
| Shakti Prasad Bavoor | Verification Engineer | NM Applications | Execution |

## Capital Equipment

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

| Equipment Type | # Equipment |
| --- | --- |
| HP ZBook G3 | 3 |
| DELL R640 | 2 |
| HP Z4 | 3 |

#### ISP13 VAL-3 Release

| Equipment Type | # Equipment |
| --- | --- |
| HP Zbook G3 | 3 |
| HP DL 360 | 5 |
| HP Z440 | 5 |

#### AVW15 RFD Release

| Equipment Type | # Equipment |
| --- | --- |
| DELL T/R 640 | 3 |
| HP DL-360 Gen9 (Rack) | 5 |
| HP ML-350 Gen9 (Tower | 2 |
| HP Z440 | 3 |
| HP Z4 (W-2145 Processor) | 3 |
| HP Z4 (W-2245 Processor) | 1 |
| HP ZBook G3 | 1 |

#### AVW15 RC Release

| Equipment Type | # Equipment |
| --- | --- |
| DELL T/R 640 | 3 |
| HP DL-360 Gen9 (Rack) | 5 |
| HP ML-350 Gen9 (Tower | 2 |
| HP Z440 | 3 |
| HP Z4 (W-2145 Processor) | 3 |
| HP Z4 (W-2245 Processor) | 1 |
| HP ZBook G3 | 1 |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

| Equipment Type | # Equipment |
| --- | --- |
| DELL T/R 640 | N/A, Execution planned on VM |
| HP DL-360 Gen9 (Rack) |  |
| HP ML-350 Gen9 (Tower |  |
| HP Z440 |  |
| HP Z4 (W-2145 Processor) |  |
| HP Z4 (W-2245 Processor) |  |
| HP ZBook G3 |  |

#### RC Release

| Equipment Type | # Equipment |
| --- | --- |
| DELL T/R 640 | N/A, Execution planned on VM |
| HP DL-360 Gen9 (Rack) |  |
| HP ML-350 Gen9 (Tower |  |
| HP Z440 |  |
| HP Z4 (W-2145 Processor) |  |
| HP Z4 (W-2245 Processor) |  |
| HP ZBook G3 |  |

### NM Suite

#### RFD Release

| Equipment Type | # Equipment |
| --- | --- |
| DELL T/R 640 | 1 |
| HP DL-360 Gen9 (Rack) | 1 |
| HP Z440 | 1 |

#### RC Release

| Equipment Type | # Equipment |
| --- | --- |
| DELL T/R 640 | 1 |
| HP DL-360 Gen9 (Rack) | 1 |
| HP Z440 | 1 |

## Key Personnel

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

| Function | Name |
| --- | --- |
| Project Verification Lead | Aashish Kumar |
| Project Core Team Lead | Daniel Broiman |
| Project Manager / RTE | Raghavendra Prasad Halaharvi |
| Quality Engineer | Shipra Shivangi |
| System Architect | Chockalingam Kavirayar |
| System Engineer | Byamma Munichandrappa |
| Lead: Regulatory Affairs Q&R | Afek Keren |
| Validation Lead | Hagit Davidesco |
| Usability Engineer | Hadar Shtern |
| PSO | Roman Grekin |
| Design Transfer | Adi Ron |

#### ISP13 VAL-3 Release

| Function | Name |
| --- | --- |
| Project Verification Lead | Aashish Kumar |
| Project Core Team Lead | Mustaq Mohammed |
| Project Manager / RTE | Raghavendra Prasad Halaharvi |
| Quality Engineer | Shipra Shivangi |
| System Architect | Chockalingam Kavirayar |
| System Engineer | Byamma Munichandrappa |
| Lead: Regulatory Affairs Q&R | Afek Keren |
| Validation Lead | Hagit Davidesco |
| Usability Engineer | Hadar Shtern |
| PSO | Roman Grekin |
| Design Transfer | Adi Ron |

#### AVW15 RFD Release

| Function | Name |
| --- | --- |
| Project Verification Lead | Aashish Kumar |
| Project Core Team Lead | Mustaq Mohammed |
| Project Manager / RTE | Raghavendra Prasad Halaharvi |
| Quality Engineer | Shipra Shivangi |
| System Architect | Chockalingam Kavirayar |
| System Engineer | Byamma Munichandrappa |
| Lead: Regulatory Affairs Q&R | Afek Keren |
| Validation Lead | Hagit Davidesco |
| Usability Engineer | Hadar Shtern |
| PSO | Roman Grekin |
| Design Transfer | Adi Ron |

#### AVW15 RC Release

| Function | Name |
| --- | --- |
| Project Verification Lead | Santhosh H |
| Project Core Team Lead | Mustaq Mohammed |
| Project Manager / RTE | Raghavendra Prasad Halaharvi |
| Quality Engineer | Shipra Shivangi |
| System Architect | Chockalingam Kavirayar |
| System Engineer | Natalia Kogan |
| Lead: Regulatory Affairs Q&R | Afek Keren |
| Validation Lead | Hagit Davidesco |
| Usability Engineer | Hadar Shtern |
| PSO | Roman Grekin |
| Design Transfer | Adi Ron |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

| Function | Name |
| --- | --- |
| Project Verification Lead | Aashish Kumar |
| Project Core Team Lead | Mustaq Mohammed |
| Project Manager / RTE | Raghavendra Prasad Halaharvi |
| Quality Engineer | Shipra Shivangi |
| System Architect | Chockalingam Kavirayar |
| System Engineer | Byamma Munichandrappa |
| Lead: Regulatory Affairs Q&R | Afek Keren |
| Validation Lead | Hagit Davidesco |
| Usability Engineer | Hadar Shtern |
| PSO | Roman Grekin |
| Design Transfer | Adi Ron |

#### RC Release

| Function | Name |
| --- | --- |
| Project Verification Lead | Santhosh H |
| Project Core Team Lead | Mustaq Mohammed |
| Project Manager / RTE | Raghavendra Prasad Halaharvi |
| Quality Engineer | Shipra Shivangi |
| System Architect | Chockalingam Kavirayar |
| System Engineer | Natalia Kogan |
| Lead: Regulatory Affairs Q&R | Afek Keren |
| Validation Lead | Hagit Davidesco |
| Usability Engineer | Hadar Shtern |
| PSO | Roman Grekin |
| Design Transfer | Adi Ron |

### NM Suite

#### RFD Release

| Function | Name |
| --- | --- |
| Project Verification Lead | Aashish Kumar |
| Project Core Team Lead | Mustaq Mohammed |
| Project Manager / RTE | Raghavendra Prasad Halaharvi |
| Quality Engineer | Shipra Shivangi |
| System Architect | Chockalingam Kavirayar |
| System Engineer | Byamma Munichandrappa |
| Lead: Regulatory Affairs Q&R | Afek Keren |
| Validation Lead | Hagit Davidesco |
| Usability Engineer | Hadar Shtern |
| PSO | Roman Grekin |
| Design Transfer | Adi Ron |

#### RC Release

| Function | Name |
| --- | --- |
| Project Verification Lead | Santhosh H |
| Project Core Team Lead | Mustaq Mohammed |
| Project Manager / RTE | Raghavendra Prasad Halaharvi |
| Quality Engineer | Shipra Shivangi |
| System Architect | Chockalingam Kavirayar |
| System Engineer | Natalia Kogan |
| Lead: Regulatory Affairs Q&R | Afek Keren |
| Validation Lead | Hagit Davidesco |
| Usability Engineer | Hadar Shtern |
| PSO | Roman Grekin |
| Design Transfer | Adi Ron |

## Estimates

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

| Sub team Name | ISP Verification |
| --- | --- |
| Dates for verification | Start Date: 03-Feb-2022 End Date: 05-Apr-2022 |
| Assumptions | None |
| Link to estimation sheet and resource planning sheet | Refer to Appendix 12.2.2 |
| Verification |  |
| Total resources for Detailed verification | 12 |
| Total working days | 38 |
| Actual Available effort | 12 * 44 = 528 |
| Total needed effort for Detailed Verification | 326 |
| Gap in effort | 0 Days |
| Gap in Resources | 0 |

#### ISP13 VAL-3 Release

| Sub team Name | ISP Verification |
| --- | --- |
| Dates for verification | Start Date: 01-Jul-2022 End Date: 12-Aug-2022 |
| Assumptions | None |
| Link to estimation sheet and resource planning sheet | Refer to Appendix 12.2.2 |
| Verification |  |
| Total resources for Detailed verification | 26 |
| Total working days | 31 |
| Actual Available effort | 27 * 31 = 837 |
| Total needed effort for Detailed Verification | 824 |
| Gap in effort | 0 Days |
| Gap in Resources | 0 |

#### AVW15 RFD Release

| Sub team Name | ISP Verification |
| --- | --- |
| Dates for verification | Start Date: 15-Sep-2022 End Date: 01-Dec-2022 |
| Assumptions | None |
| Link to estimation sheet and resource planning sheet | Refer to Appendix 12.2.2 |
| Verification |  |
| Total resources for Detailed verification | 35 |
| Total working days | 50 |
| Actual Available effort | 35 * 50 = 1550 Person Days |
| Total needed effort for Detailed Verification | 1042 Person Days |
| Gap in effort | 0 Days |
| Gap in Resources | 0 |

#### AVW15 RC Release

| Sub team Name | ISP Verification |
| --- | --- |
| Dates for verification | Start Date: 23-Mar-2023 End Date: 1-June-2023 |
| Assumptions | None |
| Link to estimation sheet and resource planning sheet | Refer to Appendix 12.2.2 |
| Verification |  |
| Total resources for Detailed verification | 20 |
| Total working days | 49 |
| Actual Available effort | 750 Person Days Note: 20*49=931, considering leaves and holiday (80% total capacity) |
| Total needed effort for Detailed Verification | 746 Person Days |
| Gap in effort | 0 Days |
| Gap in Resources | 0 |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

| Sub team Name | ISP Verification |
| --- | --- |
| Dates for verification | Start Date: 15-Sep-2022 End Date: 01-Dec-2022 |
| Assumptions | None |
| Link to estimation sheet and resource planning sheet | Refer to Appendix 12.2.2 |
| Verification |  |
| Total resources for Detailed verification | 1 |
| Total working days | 5 |
| Actual Available effort | 1 * 5 = 5 Person Days |
| Total needed effort for Detailed Verification | 2 Person Days |
| Gap in effort | 0 Days |
| Gap in Resources | 0 |

#### RC Release

| Sub team Name | ISP Verification |
| --- | --- |
| Dates for verification | Start Date: 15-Mar-2023 End Date: 01-Jun-2023 |
| Assumptions | None |
| Link to estimation sheet and resource planning sheet | Refer to Appendix 12.2.2 |
| Verification |  |
| Total resources for Detailed verification | 1 |
| Total working days | 3 |
| Actual Available effort | 1 * 3 = 3 Person Days |
| Total needed effort for Detailed Verification | 3 Person Days |
| Gap in effort | 0 Days |
| Gap in Resources | 0 |

### NM Suite

#### RFD Release

| Sub team Name | ISP Verification |
| --- | --- |
| Dates for verification | Start Date: 15-Sep-2022 End Date: 01-Dec-2022 |
| Assumptions | None |
| Link to estimation sheet and resource planning sheet | Refer to Appendix 12.2.2 |
| Verification |  |
| Total resources for Detailed verification | 2 |
| Total working days | 10 |
| Actual Available effort | 2 * 10 = 20 Person Days |
| Total needed effort for Detailed Verification | 16 |
| Gap in effort | 0 Days |
| Gap in Resources | 0 |

#### RC Release

| Sub team Name | ISP Verification |
| --- | --- |
| Dates for verification | Start Date: 16-Mar-2023 End Date: 1-June-2023 |
| Assumptions | None |
| Link to estimation sheet and resource planning sheet | Refer to Appendix 12.2.2 |
| Verification |  |
| Total resources for Detailed verification | 2 |
| Total working days | 5 |
| Actual Available effort | 2 * 5 = 10 Person Days |
| Total needed effort for Detailed Verification | 10 |
| Gap in effort | 0 Days |
| Gap in Resources | 0 |

# Verification Samples

Full requirement coverage will be performed during the verification. Each test cases will be executed at-least on one configuration.

For Software product: The sample size is always one because the code execution does not change across units and code execution is constant for the same type of hardware.

# Entry Criteria

- All the features/RFCs, as per the release scope are implemented.

- All defects planned for SVER are resolved.

- All open engineering anomalies in TFS are closed OR tagged for future milestone OR migrated as product defects.

- Design Verification Plan, Test specs for all new apps/features & RFCs are reviewed and approved.

- All tools used are validated and tool list is updated.

- All planned automation tests have corresponding validation reports reviewed and approved.

- All required documents from development (including URS, PRS, DRS, DDS, RMM, Code review reports) are reviewed and approved and Design review is complete with all identified actions closed.

- Sanity test is complete on the verification build for SVER with no open engineering anomalies to start verification.

# Exit Criteria

- All defects found during verification for the release scope have been scored and accepted.

Exception: Defects submitted during the last two weeks prior to the milestone should be set as 
     hazard yes/no (scoring is required only for Safety = Yes).

- All rank 2 defects for the release scope must be resolved and closed by end of verification.

Exceptions:

- Rank 2 defects replicated from IB projects might be postponed to the next milestone but shall be          addressed within the project.

- Rank 2 defects which need a fix outside the organization (example: US, MR etc).

- Rank 2 defects logged from outside R&D.

- All hazard / safety test cases executed for the milestone are in passed state.

- All Rank 1 defects and Sev-3 are resolved and closed.

- All the resolved defects for the release scope are closed.

- Sanity test is completed and passed on the Final Verification build during, end of verification.

- All tests including the Expert tests have been performed as per the Verification plan.

- All the documents needed for the end of verification (including Verification Reports, Traceability Matrix, RMM Update with Test Evidence & Product Defect Status Report) are reviewed and approved.

## Suspension Criteria:

- Major project scope change(s)

- Stability and reliability of the system and/or test environment

- Equipment failure of the system and/or test environment

- Excessive defects influx – Verification will be suspended if more than 5 blocker defects are logged in the applications going for Validation.

# Acceptance Criteria

On completion of verification activities, the Global Verification Lead ensures the completeness and accuracy of the Verification Records and related defects as per the exit criteria below:

- Evidence of successful completion of the Product Verification as per plan

- Approved Verification Test Reports

- Approved Verification Traceability Matrix

- Approved Product defect status report as per ICAP-P-030009 Defect Management Procedure

# Defect Management

All defects found during the verification will be logged to TFS. ICAP-P-030009 Defect Management Procedure will be followed.

# Record history

| Rev | Release Date | Author | Description of changes |
| --- | --- | --- | --- |
| A | In Windchill | Aashish Kumar | First revision of Verification Plan is created to plan the verification activities of ISP13 VAL-2 Release. - Updated all sections of this document with the changes required for ISP13 VAL-2 Release. |
| B | In Windchill | Aashish Kumar | Updated to plan the execution of RMM tests for non-impacted applications during Development Phase of ISP13 VAL-3 release. This revision is an extension of VAL-2 Verification Plan. - Updated the relevant sections – 1.1, 1.2, 1.3, 3.1. 3.3.1, 3.6, 4 and 10. |
| C | In Windchill | Aashish Kumar | A new revision of Verification Plan is created to plan the verification activities of ISP13 VAL-3 Release. - Updated all sections of this document with the changes required for ISP13 VAL-3 Release. |
| D | In Windchill | Aashish Kumar | Updated to plan the execution of RMM tests for non-impacted applications during Development Phase of AVW15 RFD release. This revision is an extension of VAL-3 Verification Plan. - Updated the relevant sections – 1.1, 1.2, 1.3, 3.1. 3.3.1, 3.6, 4 and 10. |
| E | In Windchill | Aashish Kumar | A new revision of Verification Plan is created to plan the verification activities of AVW15, CT ASPECTS & NM Suite RFD Release. - Updated all sections of this document with the changes required for of AVW15, CT ASPECTS & NM Suite RFD Release. |
| F | In Windchill | Santhosh H | A new revision of Verification Plan is created, based on the change in Test Strategy for AVW15 RFD Verification. - Updated Scope Section by creating a sub section 1.2.1.3.3.1 & 1.2.1.3.9.1. - Removed “PIM 15.0” version of VuePACS in section 1.2.1.3.14 as this version has not been released. - Updated the version of reference documents in sections 1.3.1.3, 1.3.2.1, 1.3.3.1 & 3.6.1.3. - Removed the reference of 1 document - D001157485 in section 3.3.2.2.1. - Removed the reference of 2 documents - D001059457 & D001059463 in section 1.3.1.3 & 3.6.1.3 - Removed the reference of 2 documents - D001157485 & D001157473 in section 1.3.2.1. - Added the reference of 1 document - D000993523 in section 1.3.2.1. - Removed CT Aspects Expert testing reference from section 12.4.1.3 and added in section 12.4.2. - Removed “CT Liver”, “PRVR” & “Discrete Result Export” scope of Expert testing from section 12.4.1.3. - Updated the scheduled date of Expert Testing in section 12.4.1.3. - Added KPIs in section 12.1. - Removed “Allura 8.x.100 – legacy” tests from the scope of Compatibility tests in section 1.2.1.3.14 as this version is not released. |
| G | In Windchill | Santhosh H | A new revision of Verification Plan is created to plan the verification activities of AVW15, CT ASPECTS & NM Suite RC Release. Updated all sections of this document with the changes required for of AVW15, CT ASPECTS & NM Suite RC Release. |
| H | In Windchill | Santhosh H | A new revision of Verification Plan is created to update: - Document reference for CT ASPECTS DRS and RMM in Section 1.3.2 - LNA-LCO end to end integration and claim is out of Scope and removed from section 1.2.1.4.10 |

# Record Review and Approval

| Name Author | Function Author |
| --- | --- |
| Santhosh H | Verification Lead |

| Name Reviewer(s) | Function Reviewer(s) | Review Date(s) DD-MMM-YYYY | Signature(s) |
| --- | --- | --- | --- |
| Aashish Kumar | Verification Manager | In Windchill | In Windchill |
| Chockalingam Kavirayar | System Architect | In Windchill | In Windchill |
| Natalia Kogan | System Engineer | In Windchill | In Windchill |
| Sujesh P Paravatty | Clinical Lead | In Windchill | In Windchill |
| J, Vivek Kumar | Quality Engineer | In Windchill | In Windchill |
| Adi Ron | Design Transfer | In Windchill | In Windchill |

| Name Approver | Function Approver | Date DD-MMM-YYYY | Signature |
| --- | --- | --- | --- |
| Srinivas Arakere | Senior Manager | In Windchill | In Windchill |

# Appendices

## Management of Verification Activities

| KPI Title | Description | Leading / Lagging | Target |
| --- | --- | --- | --- |
| Functional correctness | No of Features without defect/Total no features developed | Leading | >= 85% |
| Test Function completeness | Coverage by testing – traceability | Leading | = 100% |
| Performance measurement | Meeting performance specs | Leading | = 100% |
| Reliability (availability) uptime | Reliability specs | Leading | Refer to Reliability Plan (D000933496) |
| Defect density | Number of defects per KLOC | Lagging | <= 0.1 |
| No of open defect at RfD | Number of open defects | Lagging | <= 100 |

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

- Schedules – See section 4.4.

- Roles and Responsibilities – See section 4.3.

- Status Reporting – All test execution is done in TFS therefore immediate progress status is available through dashboards.

- Metrics – The following metrics will be collected in the end of each milestone:

| Metric Name | Target |
| --- | --- |
| Invalid defects found in the teams | <= 10% |

- Risk Management – See section 3.6 for product risk.
No project risk identified. All the risks are taken care in Project/Program level.

#### ISP13 VAL-3 Release

- Schedules – See section 4.4.

- Roles and Responsibilities – See section 4.3.

- Status Reporting – All test execution is done in TFS therefore immediate progress status is available through dashboards.

- Metrics – The following metrics will be collected in the end of each milestone:

| Metric Name | Target |
| --- | --- |
| Invalid defects found in the teams | <= 10% |

- Risk Management – See section 3.6 for product risk.
No project risk identified. All the risks are taken care in Project/Program level.

#### AVW15 RFD Release

- Schedules – See section 4.4.

- Roles and Responsibilities – See section 4.3.

- Status Reporting – All test execution is done in TFS therefore immediate progress status is available through dashboards.

- Metrics – The following metrics will be collected in the end of each milestone:

| Metric Name | Target |
| --- | --- |
| Invalid defects found in the teams | <= 10% |

- Risk Management – See section 3.6 for product risk.
No project risk identified. All the risks are taken care in Project/Program level.

#### AVW15 RC Release

- Schedules – See section 4.4.

- Roles and Responsibilities – See section 4.3.

- Status Reporting – All test execution is done in TFS therefore immediate progress status is available through dashboards.

- Metrics – The following metrics will be collected in the end of each milestone:

| Metric Name | Target |
| --- | --- |
| Invalid defects found in the teams | <= 10% |

- Risk Management – See section 3.6 for product risk.
No project risk identified. All the risks are taken care in Project/Program level.

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

- Schedules – See section 4.4.

- Roles and Responsibilities – See section 4.3.

- Status Reporting – All test execution is done in TFS therefore immediate progress status is available through dashboards.

- Metrics – The following metrics will be collected in the end of each milestone:

| Metric Name | Target |
| --- | --- |
| Invalid defects found in the teams | <= 10% |

Risk Management – See section 3.6 for product risk.
No project risk identified. All the risks are taken care in Project/Program level.

#### RC Release

- Schedules – See section 4.4.

- Roles and Responsibilities – See section 4.3.

- Status Reporting – All test execution is done in TFS therefore immediate progress status is available through dashboards.

- Metrics – The following metrics will be collected in the end of each milestone:

| Metric Name | Target |
| --- | --- |
| Invalid defects found in the teams | <= 10% |

Risk Management – See section 3.6 for product risk.
No project risk identified. All the risks are taken care in Project/Program level.

### NM Suite

#### RFD Release

- Schedules – See section 4.4.

- Roles and Responsibilities – See section 4.3.

- Status Reporting – All test execution is done in TFS therefore immediate progress status is available through dashboards.

- Metrics – The following metrics will be collected in the end of each milestone:

| Metric Name | Target |
| --- | --- |
| Invalid defects found in the teams | <= 10% |

Risk Management – See section 3.6 for product risk.
No project risk identified. All the risks are taken care in Project/Program level.

#### RC Release

- Schedules – See section 4.4.

- Roles and Responsibilities – See section 4.3.

- Status Reporting – All test execution is done in TFS therefore immediate progress status is available through dashboards.

- Metrics – The following metrics will be collected in the end of each milestone:

| Metric Name | Target |
| --- | --- |
| Invalid defects found in the teams | <= 10% |

Risk Management – See section 3.6 for product risk.
No project risk identified. All the risks are taken care in Project/Program level.

## Verification Table (V-Table)

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

##### Upgrade Paths – NA, the scope is only to test the fresh installation on Demo Laptop & HX deployment.

##### Estimation & Resource Planning

| Test Scenario | Config – Refer 3.2 / Verification Method | Estimation (in Days) |
| --- | --- | --- |
| MR Cardiac Suite: - Viewer - Functional LV & RV - Qflow - Spatial Enhancement - Functional LA | Full Functionality + RMM Tests | 207 |
| CT Application – CT Liver | Full Functionality Tests | 57 |
| CT Application – CCA | Full Functionality + RMM Tests | 57 |
| Installation | Sanity Tests | 4 |
| Anti-Virus Scan | N/A | 1 |
| Total Effort |  | 326 |

#### ISP13 VAL-3 Release

##### Upgrade Paths – NA, the scope is only to test the fresh installation on Demo Laptop, PIAB & HX deployment.

##### Estimation & Resource Planning

| Test Scenario | Config – Refer 3.2 / Verification Method | Estimation (in Days) |
| --- | --- | --- |
| MR Cardiac Suite: - Temporal Analysis - Functional LV RV - Mapping | Full Functionality + RMM Tests | 100 |
| LOBI CVS | Full Functionality + RMM Tests | 20 |
| MR Cardiac Suite | RMM + Regression | 40 |
| MR Apps | RMM + Regression | 20 |
| NM Apps | RMM + Sanity | 10 |
| CAD RADS | Regression | 20 |
| CCA & TAVI Legacy | RMM | 20 |
| AutoCS | Full Functionality + RMM + Regression | 25 |
| sAVA | Full Functionality + RMM + Regression | 25 |
| AVA regression | Regression | 35 |
| Trauma Legacy RMF | RMM | 15 |
| sMGOP/LMG 3D and MPR | Full Functionality + RMM | 100 |
| sCTV RFCs | Full Functionality + RMM | 50 |
| Legacy CT Apps | RMM | 25 |
| Legacy MR Apps | RMM | 25 |
| Legacy MR Cardiac Apps | RMM | 25 |
| Legacy NM Apps | RMM | 25 |
| NLog | Full Functionality | 4 |
| System feature regression | Regression | 26 |
| System RMM | Regression | 20 |
| PMCF for MR cardiac | Full Functionality | 1 |
| Labeling | Regression | 1 |
| Installation & Upgrades | Regression | 15 |
| Sanity on Portal Verification ISOs | Sanity | 50 |
| F2As Verification | Exploratory | 50 |
| Exploratory Testing | Exploratory | 50 |
| Documentation | N/A | 15 |
| Expert Testing | N/A | 11 |
| Anti-Virus Scan | N/A | 1 |
| Total Effort |  | 824 |

#### AVW RFD Release

##### Upgrade Paths

| Sl. No | Base version | Installation Path | User | Upgrade Type-Server | HW/VM | Multi Server | Formal |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | ISP11.1.8.11060 | Default | FSE | C Format | VM | Yes | Yes |
| 2 | ISP11.1.7.11077 | Default | FSE | C&D Format | HW |  |  |
| 3 | ISP12.1.5.11492 (FSE) | Default | Default | C&D Format | HW | Yes |  |
| 4 | ISP12.1.6.11044 (FSE) | Default | Default | Quick | VM |  | Yes |
| 5 | ISP12.1.6.11044 (Domain User) | Default | Domain | C Format | HW |  | Yes |
| 6 | ISP12.1.5(Domain User) | Default | Domain | Quick | VM | Yes |  |
| 7 | ISP11.1.7.11077 | Default | Default | C Format | IX |  | Yes |
| 8 | ISP12.1.5.11492 | Default | Default | Quick | IX |  | Yes |
| 9 | ISP12.1.6.11044 | Default | Default | C &D Format | Concerto |  |  |
| 10 | ISP12.1.5.11492 (Domain) | Default | Domain | C Format | Concerto |  | Yes |

##### Estimation & Resource Planning

| Test Scenario | Config – Refer 3.2 / Verification Method | Estimation (in Days) |
| --- | --- | --- |
| Findings, Bookmarks & Save Functionality | Full Functionality + Regression | 7 |
| RMM Tests | RMF | 10 |
| Regression Tests | Regression | 80 |
| Localization Tests | Regression | 15 |
| Performance Tests | Regression | 3 |
| Standard Reporting | Full Functionality | 5 |
| SAVA system requirements | Regression | 5 |
| CT Apps compatibility tests | Regression | 5 |
| CT Apps PRS tests | Regression | 5 |
| CT Apps Localization | Localization | 10 |
| Dual Energy RMF | RMF | 7 |
| BP RMF | RMF | 5 |
| CS RMF | RMF | 7 |
| 3D modeling RMF | RMF | 5 |
| BMD RMF | RMF | 7 |
| VC RMF | RMF | 10 |
| LNA RMF | RMF | 10 |
| Rebranding - Application name | Full Functionality | 3 |
| Rebranding - Labeling | Full Functionality + Regression | 3 |
| Upgrade Scenarios | Regression | 5 |
| PMCF tests | Regression | 2 |
| ZFP content | RMF | 4 |
| 3rd Party application | Full Functionality | 2 |
| System Features HFA PRS tests | Regression | 40 |
| System Localization | Localization | 5 |
| Cycle 2 | Sanity | 10 |
| CAD RADS | Full Functionality + Regression | 5 |
| Findings Management | Full Functionality + Regression | 5 |
| Discrete Report | Full Functionality | 5 |
| RFCs | Full Functionality + Regression | 10 |
| RMF test Execution | RMF | 40 |
| sMGOP/LMG 3D and MPR | Full Functionality + RMM | 15 |
| sCTV RFCs | Full Functionality + RMM | 15 |
| Regression Tests | Regression | 25 |
| Localization Tests | Regression | 35 |
| RMM Tests | RMF | 100 |
| CT Viewing RMF | RMF | 5 |
| 3rd Party NM Applications | Regression | 15 |
| Legacy PVT Owned CT, NM and MR Apps RMM, Localization, confidence test | Sanity, RMM Execution | 105 |
| 3rd Party Decoupling | Full functionality + Regression | 5 |
| Client Installer Re-design | Full functionality + Regression | 5 |
| Installation, Confidence test, Security and Upgrade scenarios | Regression | 95 |
| DynaCad 5.0 Full functionality | Full functionality + Regression | 10 |
| License – KGS & EKGS | Regression | 3 |
| System Features | Regression | 30 |
| RMM Tests | RMM | 15 |
| Localization | Regression | 12 |
| Compatibility Tests | Regression | 9 |
| IOT | Full functionality | 20 |
| Serviceability Features | Regression | 20 |
| Confidence & Exploratory Tests | Exploratory | 24 |
| Performance Tests | Full Functionality | 10 |
| Confidence & Exploratory Tests | Exploratory | 63 |
| F2As Verification | Exploratory | 41 |
| Sanity Tests | Sanity | 30 |
| Total Effort |  | 1042 |

#### AVW RC Release

##### Upgrade Paths

| Sl. No | Base version | Installation Path | User | Upgrade Type-Server | HW/VM | Multi Server | Formal |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | ISP11.1.8.11060 | Default | Default | C Format | VM | Yes | Yes |
| 2 | ISP11.1.7.11077 | Default | Domain | C&D Format | HW |  |  |
| 3 | ISP11.1.8.11060 | Install Anywhere | Domain | Fresh + | VM |  |  |
| 4 | ISP12.1.6 (FSE) | Default | Domain | C Format | HW | Yes |  |
| 5 | ISP12.1.7 (FSE) | Default | Default | C&D Format | VM |  | Yes |
| 6 | ISP12.1.7 (FSE) | Default | Default | Quick | HW | Yes |  |
| 7 | ISP12.1.6 (Domain User) | Default | Domain | Quick | HW |  | Yes |
| 8 | ISP12.1.6 | Default | Default | Quick | VM (Nutanix) | Yes | Yes |
| 9 | ISP12.1.7 (Domain + Install anywhere) | Install Anywhere | Domain | C Format | VM |  | Yes |
| 10 | ISP13 RFD | Default | Default | C&D Format | HW |  |  |
| 11 | ISP13 RFD | Default | Domain | Quick Upgrade | HW |  | Yes |
| 12 | ISP11.1.8 | Default | Default | C Format | IX |  | Yes |
| 13 | ISP12.1.6 | Default | Default | C&D Format | IX |  | Yes |
| 14 | ISP12.1.7 | Default | Default | Quick | IX |  | Yes |
| 15 | ISP13 RFD | Default | Default | Quick | IX |  |  |
| 16 | ISP11.1.8 (2-Tier) | Default | Default | C&D Format | Concerto |  | Yes |
| 17 | ISP12.1.6 (3-Tier) | Default | Default | C Format | Concerto |  | Yes |
| 18 | ISP13 RFD (2-Tier) | Default | Domain | C Format | Concerto |  |  |

##### Estimation & Resource Planning

| Test Scenario | Config – Refer 3.2 / Verification Method | Estimation (in Days) |
| --- | --- | --- |
| MRC RMM | RMM | 15 |
| MRC Full Functionality Test | Full Functionality Test | 30 |
| MRC Sanity test | Sanity test | 5 |
| MR Apps RMM | RMM | 8 |
| MR Apps Sanity | Sanity Test | 2 |
| CT Apps Standard Reporting | RMF | 2 |
| CT Apps AVA Regression | Regression | 10 |
| CT Apps BP functionality | New functionality +Regression | 15 |
| CT Apps Upgrade Scenarious | Regression | 8 |
| System HFA- Regression | Regression | 20 |
| CT Apps Regression | Regression | 10 |
| Other RMF | RMF | 15 |
| CT Apps CT Pulmo RMF | RMF | 3 |
| CT Apps QR RMF | RMF | 4 |
| CT Cardiac Apps CAD-RADS and Findings Management | Regression test (Includes defect Verification) Functional test, RMF execution, Sanity and Upgrades | 20 |
| CT Cardiac Apps CCA RFCs | Regression test (Includes defect Verification) Functional test, RMF execution, Sanity and Upgrades | 15 |
| CT Cardiac Apps CAD-RADS and Findings Management | Regression test (Includes defect Verification) Functional test, RMF execution, Sanity and Upgrades | 20 |
| CT Spectral | Regression test (Includes defect Verification) Functional test, RMF execution, Sanity and Upgrades | 20 |
| CT Viewing - Angio CT | Regression test, Sanity and Upgrades | 15 |
| NM Suite Sanity Testing | Sanity Testing only | 10 |
| NM 3rd Party Apps Sanity & Regression | Sanity and Regression testing | 5 |
| NM MiradaXD Integration Sanity & Regression | Sanity and Regression testing | 10 |
| NM Knowledgescape Sanity & Regression | Sanity and Regression testing | 5 |
| Legacy PVT Owned CT, NM, and MR Apps RMM, Localization, confidence test | RMM Execution | 114 |
| Legacy PVT Owned CT, NM, and MR Apps Apps Sanity | Sanity test | 30 |
| Installation, Confidence test, and Upgrade scenarios | Regression | 106 |
| Security | Regression | 16 |
| Nutanix | Regression | 10 |
| System Features - PMT Client Upgrade | Full functionality | 7 |
| System Features - Dynacad interventional process + Dynacad 5.1 Upgrade | Full functionality + Regression | 8 |
| System Features - 3rd party de-coupling upgrade testing + NQ + Mirada Upgrade + Dynacad Upgrade Testing | Regression | 5 |
| System Features - New RabbitMQ version Gen2 and gen3 | Regression | 12 |
| System Features - KGS Testing | Regression | 3 |
| System Features - IBE Testing | Regression | 5 |
| System Features - Sanity | Sanity | 12 |
| System Features - Regression | Regression | 15 |
| System Features - F2A | Exploratory | 20 |
| System Features - Confidence Test | Exploratory | 14 |
| Serviceability IoT | Regression | 8 |
| Serviceability Axeda | Regression | 6 |
| Serviceability Split Package testing | Regression | 3 |
| Serviceability Confidence | Exploratory | 3 |
| Serviceability Sanity | Sanity | 4 |
| Serviceability-F2A | Exploratory | 8 |
| Performance Tests | Full Functionality | 10 |
| Defect Verification | Exploratory | 40 |
| Confidence Test | Exploratory | 20 |
| Total Effort |  | 746 |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

##### Upgrade Paths – N/A, will be covered as part of AVW15 upgrade paths.

##### Estimation & Resource Planning

| Test Scenario | Config – Refer 3.2 / Verification Method | Estimation (in Days) |
| --- | --- | --- |
| Basic Functionality | Regression | 2 |

#### RC Release

##### Upgrade Paths – N/A, will be covered as part of AVW15 upgrade paths.

##### Estimation & Resource Planning

| Test Scenario | Config – Refer 3.2 / Verification Method | Estimation (in Days) |
| --- | --- | --- |
| Basic Functionality | Regression | 2 |

### NM Suite

#### RFD Release

##### Upgrade Paths – N/A, will be covered as part of AVW15 upgrade paths.

##### Estimation & Resource Planning

| Test Scenario | Config – Refer 3.2 / Verification Method | Estimation (in Days) |
| --- | --- | --- |
| NM Planar Suite | RMM + Regression | 5 |
| NM JETPack | RMM + Regression | 5 |
| NM QA QC Tools | Regression | 2 |

#### RC Release

##### Upgrade Paths – N/A, will be covered as part of AVW15 upgrade paths.

##### Estimation & Resource Planning

| Test Scenario | Config – Refer 3.2 / Verification Method | Estimation (in Days) |
| --- | --- | --- |
| NM Planar Suite | Sanity | 5 |
| NM JETPack | Sanity | 5 |
| NM QA QC Tools | Sanity | 2 |

## Exploratory Testing

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

| Requirement | Description of Exploratory Test for Requirement | Rationale |
| --- | --- | --- |
| F2As Verification | Exploratory testing to be performed for impacted features of Intellispace Portal while verifying the defects. | Defect fix may impact other features. |
| Exploratory Testing | Exploratory testing to be performed on impacted features of IntelliSpace Portal using the automation scripts for Quality & Stability of ISP. | New feature development and defect fixes may impact the quality & Stability of ISP. |

#### ISP13 VAL-3 Release

| Requirement | Description of Exploratory Test for Requirement | Rationale |
| --- | --- | --- |
| F2As Verification | Exploratory testing to be performed for impacted features of Intellispace Portal while verifying the defects. | Defect fix may impact other features. |
| Exploratory Testing | Exploratory testing to be performed on impacted features of IntelliSpace Portal using the automation scripts for Quality & Stability of ISP. | New feature development and defect fixes may impact the quality & Stability of ISP. |

#### AVW RFD Release

| Requirement | Description of Exploratory Test for Requirement | Rationale |
| --- | --- | --- |
| F2As Verification | Exploratory testing to be performed for impacted features of Intellispace Portal while verifying the defects. | Defect fix may impact other features. |
| Exploratory Testing | Exploratory testing to be performed on impacted features of IntelliSpace Portal using the automation scripts for Quality & Stability of ISP. | New feature development and defect fixes may impact the quality & Stability of ISP. |

#### AVW RC Release

| Requirement | Description of Exploratory Test for Requirement | Rationale |
| --- | --- | --- |
| F2As Verification | Exploratory testing to be performed for impacted features of Intellispace Portal while verifying the defects. | Defect fix may impact other features. |
| Exploratory Testing | Exploratory testing to be performed on impacted features of IntelliSpace Portal using the automation scripts for Quality & Stability of ISP. | New feature development and defect fixes may impact the quality & Stability of ISP. |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

| Requirement | Description of Exploratory Test for Requirement | Rationale |
| --- | --- | --- |
| None |  |  |

#### RC Release

| Requirement | Description of Exploratory Test for Requirement | Rationale |
| --- | --- | --- |
| None |  |  |

### NM Suite

#### RFD Release

| Requirement | Description of Exploratory Test for Requirement | Rationale |
| --- | --- | --- |
| None |  |  |

#### RC Release

| Requirement | Description of Exploratory Test for Requirement | Rationale |
| --- | --- | --- |
| None |  |  |

## Expert Testing

### Advance Visualization Workspace 15

#### ISP13 VAL-2 Release

| Application/Feature | Expert Test to be done by | Schedule | Effort (in Person Days) |
| --- | --- | --- | --- |
| MR Cardiac Suite: - Viewer - Functional LV & RV - Qflow - Spatial Enhancement - Functional LA | Ruth Katz, K Jayaprakash & Rishov Majumder | 20-Feb-2022 to 10-Mar-2022 | 5 PD |
| CT Applications: - CT Liver | Suru Narayan Srikrishna | 20-Feb-2022 to 10-Mar-2022 | 1 PD |
| CT Application – CCA - Findings Management - CAD-RADS - Multi-Batch | Sajimon Emmanuel | 20-Feb-2022 to 10-Mar-2022 | 2 PD |
| CT Spectral Enhancements - SMGoP - PRVR | Liora Mazangia, Lior Segev Yevgeni Zeltcer | 20-Feb-2022 to 10-Mar-2022 | 1 PD |
| RFCs - CTV - CCTV - TAVI - BP | Sajimon Emmanuel Anat Feiglin | 20-Feb-2022 to 10-Mar-2022 | 1 PD |

#### ISP13 VAL-3 Release

| Application/Feature | Expert Test to be done by | Schedule | Effort (in Person Days) |
| --- | --- | --- | --- |
| MR Cardiac Suite: - Functional LV & RV - Temporal Enhancement - Mapping - Sanity and workflow - Viewer - Qflow - Spatial Enhancement Functional LA | Ruth Katz, Rishov Majumder | 17-Jul-2022 to 27-Jul-2022 | 5 PD |
| MR LOBI CVS | Rishov Majumder | 17-Jul-2022 to 22-Jul-2022 | 1PD |
| MM AVA Spectral AVA | Shuli Ganashvili | 17-Jul-2022 to 22-Jul-2022 | 2PD |
| CT Applications: Auto CS | Suru Narayan Srikrishna | 17-Jul-2022 to 22-Jul-2022 | 1 PD |
| CT Spectral Enhancements - SMGoP PRVR | Chen Strimling, Yevgeni Zeltcer | 17-Jul-2022 to 22-Jul-2022 | 1 PD |
| RFCs - CTV - CCTV - TAVI BP | Sajimon Emmanuel Anat Feiglin | 17-Jul-2022 to 22-Jul-2022 | 1 PD |

#### AVW15 RFD Release

| Application/Feature | Expert Test to be done by | Schedule | Effort (in Person Days) |
| --- | --- | --- | --- |
| CT CCA – CAD-RADS, Findings Management, Multi-Batch | Sajimon Emmanuel | 24-Oct-2022 to 11-Nov-2022 | 1 Day |
| CMR Suite - Viewer | Ruth Katz | 02-Oct-2022 to 12-Oct-2022 | 0.5 Day |
| CMR Suite – Result and workflow | Ruth Katz | 02-Oct-2022 to 12-Oct-2022 | 0.5 Day |
| CMR Suite – LA LV Analysis | Rishov Majumder | 02-Oct-2022 to 12-Oct-2022 | 0.5 Day |
| CMR Suite – LVRV Analysis | Ruth Katz | 02-Oct-2022 to 12-Oct-2022 | 1 Day |
| CMR Suite – Spatial Enhancement | Rishov Majumder | 02-Oct-2022 to 12-Oct-2022 | 0.5 Day |
| CMR Suite – Temporal Enhancement | Rishov Majumder | 02-Oct-2022 to 12-Oct-2022 | 1 Day |
| CMR Suite - Mapping | Ruth Katz | 02-Oct-2022 to 12-Oct-2022 | 1 Day |
| Qflow – CMR Suite & Standalone | Rishov Majumder | 02-Oct-2022 to 12-Oct-2022 | 1 Day |
| CT Spectral AVA | Shuli Ganashvili | 02-Oct-2022 to 07-Oct-2022 | 1 Day |
| CT Auto CS | Suru Narayan Srikrishna | 02-Oct-2022 to 07-Oct-2022 | 0.5 Day |
| MR LOBI - CVS | Rishov Majumder | 02-Oct-2022 to 07-Oct-2022 | 0.5 Day |
| Spectral Enhancement – sMGoP | Chen Strimling, Yevgeni Zeltcer | 02-Oct-2022 to 07-Oct-2022 | 0.5 Day |
| RFCs | All | 02-Oct-2022 to 12-Oct-2022 | 1 Day |

#### AVW15 RC Release

| Application/Feature | Expert Test to be done by | Schedule | Effort (in Person Days) |
| --- | --- | --- | --- |
| MR Cardiac Suite – MM Viewer | Katz, Ruth/Majumder, Rishov | 7-Apr-2023 to 14-Apr-2023 | 1 |
| MR Cardiac Suite – Results & Workflow | Katz, Ruth/Majumder, Rishov | 7-Apr-2023 to 14-Apr-2023 | 1 |
| MR Cardiac Suite – Functional LA LV | Majumder, Rishov | 7-Apr-2023 to 14-Apr-2023 | 0.5 |
| MR Cardiac Suite – Functional LV RV | Katz, Ruth/Majumder, Rishov | 7-Apr-2023 to 14-Apr-2023 | 1 |
| MR Cardiac Suite – Spatial Enhancement | Majumder, Rishov | 7-Apr-2023 to 14-Apr-2023 | 0.5 |
| MR Cardiac Suite – Temporal Enhancement | Majumder, Rishov | 7-Apr-2023 to 14-Apr-2023 | 1 |
| MR Cardiac Suite – T1, T2, T2* Mapping | Katz, Ruth/Majumder, Rishov | 7-Apr-2023 to 14-Apr-2023 | 1 |
| MR Cardiac Suite – QFlow | Majumder, Rishov | 7-Apr-2023 to 14-Apr-2023 | 0.5 |
| CT CCA – Findings management, CAD-RADS, Multi-Batch | Ganashvili, Shuli | 7-Apr-2023 to 14-Apr-2023 | 1 |
| CT Auto CS | Suru, Narayan Srikrishna | 7-Apr-2023 to 14-Apr-2023 | 0.5 |
| LOBI – CVS (FLAIR*) | Majumder, Rishov | 7-Apr-2023 to 14-Apr-2023 | 0.5 |
| Spectral Enhancements – PRVR & sMGoP | Katz, Ruth | 7-Apr-2023 to 14-Apr-2023 | 0.5 |
| Spectral AVA | Ganashvili, Shuli | 7-Apr-2023 to 14-Apr-2023 | 0.5 |
| Discrete Results Export | Ganashvili, Shuli, Suru, Narayan Srikrishna | 7-Apr-2023 to 14-Apr-2023 | 0.5 |

### CT ASPECT Scoring (CT ASPECTS)

#### RFD Release

| Application/Feature | Expert Test to be done by | Schedule | Effort (in Person Days) |
| --- | --- | --- | --- |
| CT ASPECTS | Shuli Ganashvili | 02-Oct-2022 to 07-Oct-2022 | 0.5 Day |

#### RC Release

| Application/Feature | Expert Test to be done by | Schedule | Effort (in Person Days) |
| --- | --- | --- | --- |
| CT ASPECTS | Shuli Ganashvili | 7-Apr-2023 to 14- Apr-2023 | 0.5 Day |

### NM Suite

#### RFD Release

| Application/Feature | Expert Test to be done by | Schedule | Effort (in Person Days) |
| --- | --- | --- | --- |
| None |  |  |  |

#### RC Release

| Application/Feature | Expert Test to be done by | Schedule | Effort (in Person Days) |
| --- | --- | --- | --- |
| None |  |  |  |

End of Document